## Revision Log
Remove this sheet. It is just for in-document template purposes
Date | Who | Comments
150701 | Julio Lidón-Simón | Added Standa, Fulling, Trinamic, OM PK244-01 parallel motors. Added RLS LMA encoders
Additions to 'Tests' sheet:
Independent axis (Y/N)
Can axis collide (Y/N)
Motor phase mH
Motor phase Ohms
R Cable (Ohms)
Lim- position (ects or st) from FAT
Lim+ position (ects or st) from FAT
Time between limits (s)
Datasheet stroke (units)
Lim- position (ects prehoming) to compare with FAT
Lim+ position (ects prehoming) to compare with FAT
150706 | Julio Lidón-Simón | Columns of Motor R and L in the 'tests' sheet were inverted. Corrected
Added McLennan 23HS-309 series
Added 'DISABLE' cable sheet
Added Disable cable tests in the 'tests' sheet
150714 | Julio Lidón-Simón | In Sheet 'Tests' there's an extra column for the ects/motor turn configured (apart from the other two calculated from theoretical 'Data' and st/ects between limits
|  | Corrected: | Phytron VSS axes resistance
|  | Added: | Phytron VSS 25.200.0.6 parallel
Phytron VSS 33.200.0.6 parallel
McLennan 23HS-309 series
150911 | Julio Lidón-Simón | Motors sheet: Removed NVOLT and IVOLT
Motors sheet: Corrected Nanotech R and L values
Motors sheet: PKP268MD28x-x inductance updated
Added TANGO sheet
Sheet 'System info' changed to 'SystemInfo'.
Changes to new 'SystemInfo'. Added number of cables
Data sheet: Full change of structure
Added POSITIONS sheet
Added DISABLE sheet
Tests sheet: Renamed and added a big number of columns
List of steps to update to this version:
Update motor PKP268MD28x-x
+ Modify 'Data'
- Unhide hidden data columns (E) in 'Data' sheet
- Add two new columns in A and B shifting right everything.
- Copy in the empty lines at the end of your data in 'Data' the info from columns A-D in 'Tests'. The information in 'Data' will have that structure in the new layout.
- In the rest of the fields for those lines you just fill cols A-D fill the information from the lines already present in the 'Data' sheet (take the info from its equipment type).
- Rename the sheet 'System info' to 'SystemInfo'
- Once that is done, write the number of cables in SystemInfo!$B$1 from Data!$C$2 (initially was in A2 but was moved there when the two columns were added in step 2).
- Using Replace Within:Workbook, Look in: Formulas Find what: Data!$C$2, Replace with: SystemInfo!$B$1 update all the formulas so that Data!$C$1 is taken instead
- Remove all the rows in 'Data' between the first row and the row with the column titles (the column titles will be now in row 2).
- In the sheet 'Data' correct the formulas in the columns N-Q (taking info from 'MOTORS') so that they have NO spaces in the formulas :o (dirty one).
- Using Replace Within:Workbook, Look in: Formulas Find what: $1;0, Replace with: $2;0 update all the formulas (we're going to remove the first row from 'Data' too).
- Using Replace Within:Workbook, Look in: Formulas Find what:'MATCH(CONCATENATE($C' , Replace with: 'MATCH(CONCATENATE($B'
- At the end, Add column F (sardana name) and fill it with the scientist and column AE (TC tango device)
+ Create 'POSITIONS' sheet
- Fill columns till E
- Get columns F-G from the formulas
- Fill Columns I-N from the info in Data (just copy)
- Fill Columns O and P with the info you get from your scientist (these values will be used to calculate the soft limits in sardana from the hard limits position info you get during testing).
+ Modify 'Tests' sheet
- Add E column (Sardana axis name) (takes info from Data)
- Add columns S(I*R) to V(IVOLT configured).
- Be careful here in case you have put info in old AA (Homing position) by hand during your tests. Homing columns Z and AA have to be changed by new AD (Home origin)-AH(Home done). 'Fat home position' is the info for the fat. AF contains the ABSOFFSET param you configure. AG Home position has to contain the HOMEPOS parameter you configure in the home tab of icepapcms.
This aims to separate ABSOFFSET from the fat home position.
- Inserts columns BC (stroke from encoder counts)-BO(Soft lim-). This is information for sardana configuration. Sign has to be +1 or -1. Offset defaults to 0
- Insert column BQ (gets tc device from Data)
- Insert column CB with the text chain for the tango attributes
160113 | Julio Lidón-Simón | Motors sheet:
Added NVOLT and IVOLT again.
Added following motors
Oriental PK245DB
Oriental PK245MD15B
Oriental PK245M-02B unipolar 1 coil
Oriental PK266D14B
Oriental PK268D14B
Oriental PKP264MU20B-L unipolar 1 coil
161111 | Gabriela Todorescu | Data sheet:
renaming TGT encoder to 1st encoder
1st encoder source and 2nd encoder source added
Tests sheet:
2nd encoder cable column added
columns CI to CW related to 2nd encoder added
161121 | Julio Lidón Simón | Motors sheet:
Added all 3 missing motors to the motor sheet of the excel
Oriental PK268M-01 half coil
Nanotech AS4118L1804-E
Nanotech AS5918L4204-E
Added Renishaw 32bits MSBMSKNB 1 (31b)
Changed unipolar 1 coil motors to half coil
Tests sheet:
Added a column for NRES in TESTS
Data sheet:
Added two columns for repeatability and resolution in Data
161206 | Julio Lidón
- Data.Repeatabilty header changed (AN)
- Data. Disable signals (AL) changed to: Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd
- Configurations. C1 changed to: Axis name (ccdb)
- Configurations. I column formula was wrong (taken from 2 rows below than expected (from row 4 for row 2).
- FATPOSITIONS column F formula was wrong (taken from 2 rows below than expected (from row 4 for row 2).
- Tests. Column D header changed to: Axis name (ccdb)
- Tests. Column AA header changed to: NRES (?MEAS R)
- Tests. Column AF header changed to: Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd
- Tests.AY. Inserted new column with header: Encoder noise pkpk (ects)
- Tests.BR. Formula changed to have into account UseEncoderSource to: =IF(BP17="FALSE";($BM17*($AQ17/$BL17)+$BQ17);($BM17*($AR17/$AN17)+$BQ17))
- Tests.BP. Formula changed to have into account UseEncoderSource to: =IF(BP17="FALSE";($BM17*($AS17/$BL17)+$BQ17);($BM17*($AT17/$AN17)+$BQ17))
- Tests.BP column header changed to: UseEncoderSourceFormula: TRUE/FALSE
- Tests CA column: Inserted new column with header: Warning temperature
- Tests CB column: Inserted new column with header: Alarm temperature
- Tests CG column: Inserted new column with header: Encoder noise pkpk (units). Formula: =$AY2/$AN2
- Tests CH column: Inserted new column with header:Deadband in (units). Formula:=$CF2/($AP2*$AM2)
- Tests CI column: Inserted new column with header: Aimed repeatability (units). Formula:=Data!$AN2
- Tests CJ column: Inserted new column with header: Aimed resolution (units). Formula:=Data!$AO2
- Motors. Phytron ZSS.52.200.2.5 parallel appeared twice
- Motors. Oriental PK243-01-SGXX series renamed to Oriental PK243-01 series
- Motors. Oriental PK243-02-SGXX series renamed to Oriental PK243-02 series
- Motors. Oriental PK245DA is renamed to Oriental PK245D (since DA and DB are the same motor with just single or double shaft).
- Motors. Oriental PK245MD15B renamed to Oriental PK245MD15 (since A or B stand for the double shaft version)
- Motors. Oriental Oriental PK245M-02B series renamed to Oriental PK245M-02 series (final B stands for double shaft version)
- Motors. Oriental PK245M-02B half coil renamed to Oriental PK245M-02 half coil
- Motors. Oriental PK245M-02 half coil rated current increased by a factor 1.4
- Motors. Oriental PK266D14B renamed to Oriental PK266D14
- Motors. Oriental PK264JD renamed to Oriental PK264JD
- Motors. Oriental PK267JA series renamed to Oriental PK267J series
- Motors. Added Oriental PK267JD series
- Motors. Oriental PK268D14B renamed to Oriental PK268D14
- Motors. Oriental PK268M-01 half coil rated current increased by a factor 1.4
- Motors. Oriental PK296-AE-SGxx series renamed to Oriental PK296AE-SGxx series
- Motors. Oriental PK296-AE-SGxx parallel renamed to Oriental PK296AE-SGxx parallel
- Data. Added column AP for settling time constraints for fast step scans
- Tests. Added column CP for settling time constraints for fast step scans
- Tests. Added column CQ for validation text with who and when
- Manufacturers info sheet updated with document: https://alfresco.maxlab.lu.se/share/page/site/maxiv/document-details?nodeRef=workspace://SpacesStore/7adbde15-a2e8-41ef-a236-9060bda8fe38#
- Motors. NVOLT has been limited to 48V in the excel (to avoid driving motors with low resistance, low inductance and high rated currents with short test cables on too much voltage, you can increase the voltage if you need it for speed reasons yourself)
170116 | Julio Lidón | - Tests. Column CS with tango attribs chain. CONCATENATE function does not keep the number formatting and creates problems with . and , as separator for decimals. The symbol & solves that problem.
Therefore the column now has the following formula:
="Step_per_unit:"&$BL2&";"&"Offset:"&$BQ2&";"&"Sign:"&$BM2&";"&"EncoderSource:"&$BN2&";"&"EncoderSourceFormula:"&$BO2&";"&"UseEncoderSource:"&$BP2
170130 | Gabriela Todorescu
Tango sheet
"device" -> "devicename"
"controller sardana" -> "controller"
"axis name sardana" -> "name"
"axis nr" -> "axis"
"attribute chain" -> "attributes"
"units" -> "unit"
Ready. (N or blank: don't take the data here)
Instrument column added
Data sheet
column AB added; Motor mechanical resolution (full steps /2nd unit)
Tests sheet
column CY added; Motor resolution (fst per mm/deg)
column DC formula changed to =IF(K2="-";"-";(CZ2/CY2)*AO2); Theoretical 2nd encoder counts per motor turn

## Manufacturers info
S2995
Motion Axes | [11 empty] | Note: Motors wired in Parallel: Run Current = 1.4 x Rating Plate Current
Last edited | 2015-06-25 | [10 empty] | Note: Motors wired in Series: Run Current = 0.7 x Rating Plate Current
by | Adam Young
fst full step, hst half step
Component | Actuation | Motor model | Mechanical transmission details | Air/Vacuum | Step operation | Wiring (parallel/series) | Rated current (for specified wiring) | Rated idle Current | Motor mechanical resolution (meter or angle per step) | Range (fst or m) | Speed (cts/msec) | Encoder Res. (meter/count) | Encoder readhead, interpolator and scale model | Home switch | Limits | Overtravel Limits | Thermocouples for EPS / Motor | Serial Device | Camera | Mech. Drwg | Elec. Drwg

## SystemInfo
Number of rows/cables: | 50
IcePAP system master name | B316A-A101630-CAB05-CTL-IPAP-01.maxiv.lu.se
IcePAP system master mac | 00-0C-C6-79-BA-74
IcePAP system master IP | 172.16.134.41
Number of chassis | 4
Number of drivers | 32
Alfresco link to cable drawings
Alfresco link to hardware overview | https://alfresco.maxiv.lu.se/share/page/site/veritas/document-details?nodeRef=workspace://SpacesStore/084a3503-983a-4853-b86b-86f0f70c4258
Alfresco link to motorization information excel SAT | https://alfresco.maxiv.lu.se/share/page/site/veritas/document-details?nodeRef=workspace://SpacesStore/c9defc2e-3bb9-4f23-bca2-48ae0c309118
Labels printed
Client computer | b-v-veritas-cc-0.maxiv.lu.se
Machine hosting icepapcms database | b-v-veritas-csdb-0.maxiv.lu.se
Sardana pool generated | B316A/Pool/01

## Data
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Axis name (sardana) | Motor model and wiring (=template) | Independent axis (Y/N) | Can axis collide (Y/N) | Mechanical transmission details | Unit | Air/Vacuum | Step division operation | Sinusoidal peak current (A) | Resistance per phase (Ohm) | Inductance per phase(mH) | Motor full steps | Recommended idle current % of nominial | Motor mechanical resolution (full steps /unit) | Range (units) | Max speed(fst/s) | 1st encoder source (ABSENC, ENCIN, INPOS, NONE) | Theoretical 1st encoder counts per motor turn | 1st encoder Res. (counts/unit) | 1st encoder readhead, interpolator and scale model | 1st encoder template (if absolute) | Home origin (absolute, limit-, encoder, homesw) | 1st encoder supply (Voltage supply, signal levels (TTL/RS422) and current consumption) | Motor mechanical resolution (full steps /2nd unit) | 2nd encoder source (ABSENC, ENCIN, INPOS, NONE) | Theoretical 2nd encoder counts per motor turn | 2nd encoder Res. (counts/unit) | 2nd encoder readhead, interpolator and scale model | 2nd encoder template (if absolute) | 2nd encoder supply (Voltage supply, signal levels (TTL/RS422) and current consumption) | Limits ('-' if none) | Overtravel Limits ('-' if none) | Thermocouples for EPS / Motor ('-' if none) | Associated TC device | Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd | Limit switches type | Repeatability (units) | Resolution (units) | Settling time (if fast step scans needed) | Motor direction | Backlash | Parasitic effects | Serial Device | Camera | Mech. Drwg | Elec. Drwg
1 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | a_m4_baff_vt | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 30mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
2 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | a_m4_baff_vb | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 30mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
3 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | a_m4_baff_hl | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 20mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
4 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | a_m4_baff_hr | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 20mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
5 | B316A-EA03-DIA-BPM-01 | M4 BPM VERITAS | V | a_m4_bpm_v | Oriental PKP268MD28B-L | Y | N | Clean Damper D6CL-8.0F; | mm | Air | 2 | 3.9 | 1.23 | 6 | 400 | 50 | 400 | 50mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B15F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry D45X
6 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | a_m4_v1 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
7 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | a_m4_v2 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
8 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | a_m4_v3 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
11 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | a_m4_h4 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
12 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | a_m4_h5 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
13 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | X | a_mp1_x | Oriental PK245M-02B series | Y | N | 2mm pitch | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 200 | ±15mm |  | ABSENC | 2000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 |  |  | OPEN
14 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Z | a_mp1_z | Oriental PK245M-02B series | Y | N | 2mm pitch | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 200 | ±15mm |  | ABSENC | 2000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 |  |  | OPEN
15 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Y | a_mp1_y | Oriental PK264M-E2.0 parallel | Y | N | 2:1, 4mm pitch | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | 50 | 200 | 358mm(400mm) |  | ABSENC | 2000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 |  |  | OPEN
16 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | YAW | a_mp1_yaw | Nanotec L4118S1404-M6X1 | Y | N |  | deg | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 |  |  | ABSENC | 11930445.5027 | 11930445.5027 | Renishaw RA32BAA100B10A | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 |  |  | OPEN
17 | B316A-EA05-DIA-SPC-01 | Q-Chamber VERITAS | YAW | a_spc1_yaw | Oriental PK296-E4.5 parallel | Y | N |  | deg | Air | 2 | 7.5 | 0.24 | 1.5 | 200 | 50 | 200 |  |  | ABSENC | 0 |  | RE36SC0613B10A2A06 |  | absolute
18 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VT | a_gr_baff_vt | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
21 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VB | a_gr_baff_vb | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
22 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HR | a_gr_baff_hr | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
23 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HL | a_gr_baff_hl | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
24 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y1 | a_gra1_y1 | Newport 8301-UHV | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  |  | #N/A |  | Microsense, 4800 UHV + Motrona UZ210
25 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y2 | a_gra1_y2 | Newport 8301-UHV | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  |  | #N/A |  | Microsense, 4800 UHV + Motrona UZ210
26 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y3 | a_gra1_y3 | Newport 8301-UHV | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  |  | #N/A |  | Microsense, 4800 UHV + Motrona UZ210
27 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y1 | a_gra2_y1 | Newport 8301-UHV | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  |  | #N/A |  | Microsense, 4800 UHV + Motrona UZ210
28 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y2 | a_gra2_y2 | Newport 8301-UHV | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  |  | #N/A |  | Microsense, 4800 UHV + Motrona UZ210
31 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y3 | a_gra2_y3 | Newport 8301-UHV | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  |  | #N/A |  | Microsense, 4800 UHV + Motrona UZ210
32 | B316A-EA08-OPT-FLT-01 | Filter Unit VERITAS | X | a_flt1_x | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B10A, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
33 | B316A-EA08-DIA-DETMC-01 | Detector VERITAS | Y | a_detmc1_y | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
34 | B316A-EA08-DIA-DETMC-02 | Detector VERITAS | Y | a_detmc2_y | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
35 | B316A-EA09-OPT-MM-01 | Movable Mask VERITAS | Y | a_mm1_y | Oriental PK268M-E2.0B series | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  | ABSENC | #N/A | 1000000 | Renishaw RL32BAT001B10A, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
36 | B316A-EA09-OPT-MM-02 | Movable Mask VERITAS | Y | a_mm2_y | Oriental PK268M-E2.0B series | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  | ABSENC | #N/A | 1000000 | Renishaw RL32BAT001B10A, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
37 | B316A-EA09-OPT-POL-01 | Multilayer Mirror VERITAS | Y | a_pol1_y | Oriental PK264M-E 2.0B | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  | ABSENC | #N/A | 1000000 | Renishaw RL32BAT001B10A, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
38 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | PIT | a_pol2_pit | Oriental PK245M-02B series | Y | N |  | deg | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
41 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | X | a_pol2_x | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
42 | B316A-EA09-DIA-TABLED-01 | Detector Table VERITAS | Z | a_tabled_z | Oriental PK299-E4.5B parallel | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  |  | #N/A |  | Givi misure AGS | Renishaw Resolute 32 bits BISS-C | absolute
43 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | PIT | a10_detmc1_pit | Oriental PK245M-02B series | Y | N |  | deg | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 0 |  | Renishaw RA32BAA100B10A | Renishaw Resolute 32 bits BISS-C | absolute
44 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | Y | a10_detmc1_y | Oriental PK264M-E2.0B parallel | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50 | #N/A |  |  | ABSENC | #N/A | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute
45 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT1 | a_detmc2_pit1 | Oriental PK245M-02B series | Y | N |  | deg | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  | ABSENC | 0 |  | Renishaw RA32BAA199B10A | Renishaw Resolute 32 bits BISS-C | absolute
46 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT2 | a_detmc2_pit2 | Oriental PK264M-E2.0 parallel | Y | N |  | deg | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | 50 | 400 |  |  | ABSENC | 0 |  | Renishaw RA32BAA150B10A | Renishaw Resolute 32 bits BISS-C | absolute
47 | B316A-EA11-DIA-SPC-01 | Spectrometer VERITAS | YAW | a11_spc1_yaw | Oriental PK296-E4.5 parallel | Y | N |  | deg | Air | 2 | 7.5 | 0.24 | 1.5 | 200 | 50 | 200 |  |  |  | 0
48 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | X | a_mp2_x | Oriental PK268M-E2.0 series | Y | N |  | mm | Air | 2 | 2 | 4.5 | 19.2 | 400 | 50 | 400 |  |  |  | 0
51 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Z | a_mp2_z | Oriental PK245M-02B series | Y | N |  | mm | Air | 2 | 0.8 | 15 | 39.6 | 400 | 50 | 400 |  |  |  | 0
52 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Y | a_mp2_y | Oriental PK264M-E2.0 parallel | Y | N |  | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | 50 | 400 |  |  |  | 0
53 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | b_m4_baff_vt | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 30mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
1 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | b_m4_baff_vb | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 30mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
55 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | b_m4_baff_hl | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 20mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
56 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | b_m4_baff_hr | Nanotec L4118S1404-M6X1 | Y | N | No damper; 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | 50 | 200 | 20mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 0 | - | - | -
57 | B316A-EB03-DIA-BPM-01 | M4 BPM VERITAS | V | b_m4_bpm_v | Oriental PKP268MD28B-L | Y | N | Clean Damper D6CL-8.0F; | mm | Air | 2 | 3.9 | 1.23 | 6 | 400 | 50 | 400 | 50mm | - | ABSENC | 1000000 | 1000000 | Renishaw RL32BAT001B15F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry D45X
58 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | b_m4_v1 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
61 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | b_m4_v2 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
62 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | b_m4_v3 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
63 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | b_m4_h4 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
64 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | b_m4_h5 | Oriental PK245M-02B unipolar 1 coil | Y | N | Clean Damper D4CL-5.0F; Harmonic Drive 100:1 Gearbox; 1mm spindle pitch | um | Air | 2 | 0.8 | 7.5 | 9.9 | 400 | 50 | 40 | 10mm (±5mm) | - | ABSENC | 10000 | 1000 | Renishaw RL32BAT001B05F, RLTA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | 5V ± 10%; 250mA; RS422 | [7 empty] | 2 | 2 | - | - | 2O | Cherry DB1C-A1AA normally closed
65 | [8 empty] | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50
66 | [8 empty] | um | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50
67 | [8 empty] | um | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50
68 | [8 empty] | um | Air | 2 | #N/A | #N/A | #N/A | #N/A | 50

## Configurations
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Motor cable | Encoder cable ('-' if none) | ANSTEP | ANTURN | TGTENC | ABSNSTEP | ABSNTURN | EINNSTEP | EINNTURN
1 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | W032648 | W032649 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
2 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | W032650 | W032651 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
3 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | W032652 | W032653 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
4 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | W032654 | W032655 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
5 | B316A-EA03-DIA-BPM-01 | M4 BPM VERITAS | V | W032656 | W032657 | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
6 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | W032658 | W032659 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
7 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | W032660 | W032661 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
8 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | W032662 | W032663 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
11 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | W032664 | W032665 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
12 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | W032666 | W032667 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
13 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | X | W050216 | W050217 | 800 | 1 | ABSENC | 2000000 | 1 |  | 1
14 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Z | W050218 | W050219 | 800 | 1 | ABSENC | 2000000 | 1 |  | 1
15 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Y | W050220 | W050221 | 800 | 1 | ABSENC | 2000000 | 1 |  | 1
16 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | YAW | W050222 | W050223 | 400 | 1 | ABSENC | 11930445.5027 | 1 |  | 1
17 | B316A-EA05-DIA-SPC-01 | Q-Chamber VERITAS | YAW | #N/A | #N/A | 400 | 1 | ABSENC | 0 | 1 |  | 1
18 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VT | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
0
0
21 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VB | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
22 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HR | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
23 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HL | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
24 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y1 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
25 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y2 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
26 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y3 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
27 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y1 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
28 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y2 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
0
0
31 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y3 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
32 | B316A-EA08-OPT-FLT-01 | Filter Unit VERITAS | X | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
33 | B316A-EA08-DIA-DETMC-01 | Detector VERITAS | Y | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
34 | B316A-EA08-DIA-DETMC-02 | Detector VERITAS | Y | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
35 | B316A-EA09-OPT-MM-01 | Movable Mask VERITAS | Y | #N/A | #N/A | #N/A | 1 | ABSENC | #N/A | 1 |  | 1
36 | B316A-EA09-OPT-MM-02 | Movable Mask VERITAS | Y | #N/A | #N/A | #N/A | 1 | ABSENC | #N/A | 1 |  | 1
37 | B316A-EA09-OPT-POL-01 | Multilayer Mirror VERITAS | Y | #N/A | #N/A | #N/A | 1 | ABSENC | #N/A | 1 |  | 1
38 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | PIT | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
41 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | X | #N/A | #N/A | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
42 | B316A-EA09-DIA-TABLED-01 | Detector Table VERITAS | Z | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
43 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | PIT | #N/A | #N/A | 800 | 1 | ABSENC | 0 | 1 |  | 1
44 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | Y | #N/A | #N/A | #N/A | 1 | ABSENC | #N/A | 1 |  | 1
45 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT1 | #N/A | #N/A | 800 | 1 | ABSENC | 0 | 1 |  | 1
46 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT2 | #N/A | #N/A | 800 | 1 | ABSENC | 0 | 1 |  | 1
47 | B316A-EA11-DIA-SPC-01 | Spectrometer VERITAS | YAW | #N/A | #N/A | 400 | 1 | 0 |  | 1 |  | 1
48 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | X | #N/A | #N/A | 800 | 1 | 0 |  | 1 |  | 1
51 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Z | #N/A | #N/A | 800 | 1 | 0 |  | 1 |  | 1
52 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Y | #N/A | #N/A | 800 | 1 | 0 |  | 1 |  | 1
53 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | W050829 | W050830 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
1 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | W050831 | W050832 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
55 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | W050833 | W050834 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
56 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | W050835 | W050836 | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
57 | B316A-EB03-DIA-BPM-01 | M4 BPM VERITAS | V | W050837 | W050838 | 800 | 1 | ABSENC | 1000000 | 1 |  | 1
58 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | W050839 | W050840 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
61 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | W050841 | W050842 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
62 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | W050843 | W050844 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
63 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | W050845 | W050846 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
64 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | W050847 | W050848 | 800 | 1 | ABSENC | 10000 | 1 |  | 1
65 | 0 | 0 | 0 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
66 | 0 | 0 | 0 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
67 | 0 | 0 | 0 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
68 | 0 | 0 | 0 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1

## FATPOSITIONS
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Units | Encoder input | Home Nominal position (ects) | FAT Home mark (ects) | FAT LIM- (ects) | FAT LIM+ (ects) | FAT Home (units) | FAT LIM- (units) | FAT LIM+ (units) | Margin soft limit- (units) | Margin soft limit+ (units)
1 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | mm | ABSENC | 1137151156 | 1141382500 | 1137397500 | 1156335550 | 0 | -3.99 | 14.95 | - | - | Test docs W002033-03; Vertical Blade Upper
2 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | mm | ABSENC | 270025089 | 285271100 | 270296000 | 289271100 | 0 | -14.98 | 4 | - | - | Test docs W002033-03; Vertical Blade Lower
3 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | mm | ABSENC | 1499183098 | 1493769300 | 1488770800 | 1503773200 | 0 | -5 | 10 | - | - | Test docs W002033-03; Horizontal Blade LHS
4 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | mm | ABSENC | 1736309985 | 1731444100 | 1726440100 | 1741418400 | 0 | -5 | 9.97 | - | - | Test docs W002033-03; Horizontal Blade RHS
5 | B316A-EA03-DIA-BPM-01 | M4 BPM VERITAS | V | mm | ABSENC | 30452456 | - | - | - | - | - | - | - | -
6 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | um | ABSENC | 84229884 | 84229529 | 79262957 | 89556843 | 84.2295 | 79.263 | 89.5568 | - | - | 20140073_CL-RABER-Linear_Legdrive_11b
7 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | um | ABSENC | 84264739 | 84265053 | 78917342 | 89182222 | 84.2651 | 78.9173 | 89.1822 | - | - | 20140073_CL-RABER-Linear_Legdrive_18b_final
8 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | um | ABSENC | 84182075 | 84179513 | 78778910 | 89084464 | 84.1795 | 78.7789 | 89.0845 | - | - | 20140073_CL-RABER-Linear_Legdrive_35b2
0 | 0 | 0 | 0 | 0
11 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | um | ABSENC | 84515611 | 84516617 | 79454625 | 89489938 | 84.5166 | 79.4546 | 89.4899 | - | - | 20140073_CL-RABER-Linear_Legdrive_38b5
12 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | um | ABSENC | 84135906 | 84135330 | 79187845 | 89575714 | 84.1353 | 79.1878 | 89.5757 | - | - | 20140073_CL-RABER-Linear_Legdrive_41b
13 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | X | mm | ABSENC
14 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Z | mm | ABSENC
15 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Y | mm | ABSENC
16 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | YAW | deg | ABSENC
17 | B316A-EA05-DIA-SPC-01 | Q-Chamber VERITAS | YAW | deg | ABSENC
18 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VT | mm | ABSENC
0
0
21 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VB | mm | ABSENC
22 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HR | mm | ABSENC
23 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HL | mm | ABSENC
24 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y1 | mm | 0
25 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y2 | mm | 0
26 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y3 | mm | 0
27 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y1 | mm | 0
28 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y2 | mm | 0
0
0
31 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y3 | mm | 0
32 | B316A-EA08-OPT-FLT-01 | Filter Unit VERITAS | X | mm | ABSENC
33 | B316A-EA08-DIA-DETMC-01 | Detector VERITAS | Y | mm | ABSENC
34 | B316A-EA08-DIA-DETMC-02 | Detector VERITAS | Y | mm | ABSENC
35 | B316A-EA09-OPT-MM-01 | Movable Mask VERITAS | Y | mm | ABSENC
36 | B316A-EA09-OPT-MM-02 | Movable Mask VERITAS | Y | mm | ABSENC
37 | B316A-EA09-OPT-POL-01 | Multilayer Mirror VERITAS | Y | mm | ABSENC
38 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | PIT | deg | ABSENC
41 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | X | mm | ABSENC
42 | B316A-EA09-DIA-TABLED-01 | Detector Table VERITAS | Z | mm | 0
43 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | PIT | deg | ABSENC
44 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | Y | mm | ABSENC
45 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT1 | deg | ABSENC
46 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT2 | deg | ABSENC
47 | B316A-EA11-DIA-SPC-01 | Spectrometer VERITAS | YAW | deg | 0
48 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | X | mm | 0
51 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Z | mm | 0
52 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Y | mm | 0
53 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | mm | ABSENC
1 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | mm | ABSENC
55 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | mm | ABSENC
56 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | mm | ABSENC
57 | B316A-EB03-DIA-BPM-01 | M4 BPM VERITAS | V | mm | ABSENC
58 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | um | ABSENC |  | 84095873 | 79056547 | 89155784 | 84095.873 | 79056.547 | 89155.784
61 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | um | ABSENC |  | 78960638 | 73779002 | 83968939 | 78960.638 | 73779.002 | 83968.939
62 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | um | ABSENC |  | 84128024 | 79004220 | 89190308 | 84128.024 | 79004.22 | 89190.308
63 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | um | ABSENC |  | 84001007 | 78970165 | 89093371 | 84001.007 | 78970.165 | 89093.371
64 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | um | ABSENC |  | 81975505 | 76941717 | 87075481 | 81975.505 | 76941.717 | 87075.481
65 | 0 | 0 | 0 | mm | 0
66 | 0 | 0 | 0 | um | 0
67 | 0 | 0 | 0 | um | 0
68 | 0 | 0 | 0 | um | 0
Note on G1: encoder value when connected to icepap first time

## Tests
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Axis name (sardana) | Independent axis (Y/N) | Can axis collide (Y/N) | Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd ; open;gnd) | Motor cable | 1st encoder cable ('-' if none) | 2nd encoder cable ('-' if none) | Disable cable ('-' if none) | Cables labling checked | Motor template | Encoder template | Motor phase mH | Motor phase Ohms | measured mH | measured Ohms | R Cable (Ohms) | I*R | Vinductance (per 1khsts) | NVOLT | IVOLT | NCURR | ICURR | NRES (?MEAS R) | Lim- polarity (NORMAL/INVERTED) | Lim+ polarity(NORMAL/INVERTED) | Motor direction (NORMAL/INVERTED) | Encoder direction (NORMAL/INVERTED) | Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd ; open;gnd) | Disable tests | Home origin (absolute, limit-, encoder, homesw) | FAT Home position | Absolute encoder offset (ects) (0 if no absolute encoder) | Home position (steps) | Home done | Motor resolution (fst per mm/deg) | Encoder Res. (counts/mm or deg) | Motor steps | ustep factor | lim- steps (releasing switch) | lim- ects ('-' if none) (releasing switch) | lim+ steps (releasing switch) | lim+ects (releasing switch) | Lim- position (ects or st) from FAT | Lim+ position (ects or st) from FAT | Lim- position (ects prehoming) to compare with FAT | Lim+ position (ects prehoming) to compare with FAT | Encoder noise pkpk (ects) | Speed | Acctime | Time between limits (s) | Theoretical encoder counts per motor turn | encoder counts per motor turn | Encoder counts per motor turn configured | Range (units) | stroke (from motor steps) in mm or deg | stroke (from encoder counts) in mm or deg | Units | Steps per unit (as calculated from motor steps) | Steps per unit (as calculated from encoder) | Tango Step_per_unit | Tango motor Sign(+1 or -1) | EncoderSource | EncoderSourceFormula | UseEncoderSourceFormula: TRUE/FALSE | Tango motor Offset (units) | Lim- position (units) | Lim+ position (units) | FAT lim- position (units) | FAT lim+ position (units) | Soft lim- margin (units) | Soft lim+ margin (units) | TC(MPS/Motor) | Associated TC device | In vacuum motor temperature after 5min operation (wherever it applies, otherwise '-') | Warning temperature | Alarm temperature | Disable cable and conditions tested (without plc, with final plc, final setup…) | Disable config | CL time constant | CL deadband | Encoder noise pkpk (units) | Deadband in (units) | Aimed repeatability (units) | Aimed resolution (units) | CL error allowed | Start velocity | Settle window (steps) | Settle time(s) | Fast settling time required for fast scan | Measured settling time (s) | Validation (who, when) | Comments | Tango attribs chain | 2nd encoder template | 2nd Encoder direction (NORMAL/INVERTED) | FAT Home position | Home position (steps) | Home done | Motor resolution (fst per mm/deg) | 2nd encoder Res. (counts/unit) | lim- 2nd ects ('-' if none) (releasing switch) | lim+ 2nd ects (releasing switch) | Theoretical 2nd encoder counts per motor turn | 2nd encoder counts per motor turn | 2nd encoder counts per motor turn configured | Range (units) | stroke (from 2nd encoder counts) (units) | Units | Steps per unit (as calculated from 2nd encoder)
1 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | a_m4_baff_vt | Y | N | - | W032648 | W032649 |  | - |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.6 | 2.3 | 0.3 | 4.6 | 2.512 | 24 | 12 | 1.5 | 50 |  | N | N | INV | N | - | N/A | absolute | 1141382500 | 0 |  | N/A | 200 | 1000000 | 200 | 2 | -7686 | 1137151156 | 76 | 1156521622 | 1137397500 | 1156335550 | 1137151156 | 1156521622 |  | 400 | 0.15 | 19.405 | 1000000 | 998220.3556 | 1000000 | 30mm | 19.405 | 19.3705 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -1137.23 | -0.078844 | 19.2916 | -3.99 | 14.95 |  |  | - | - | - |  |  | - | NONE | 0.1 | 40 | 0 | 0.1 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020773; motor marked W002034/03; | Step_per_unit:400;Offset:-1137,23;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 1141382500 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 30mm |  | mm | 0
2 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | a_m4_baff_vb | Y | N | - | W032650 | W032651 |  | - |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.6 | 2.4 | 0.4 | 4.8 | 2.512 | 24 | 12 | 1.5 | 50 |  | N | N | INV | N | - | N/A | absolute | 285271100 | 0 |  | N/A | 200 | 1000000 | 200 | 2 | -86 | 269809070 | 7815 | 289522512 | 270296000 | 289271100 | 269809070 | 289522512 |  | 400 | 0.15 | 19.7525 | 1000000 | 998022.63 | 1000000 | 30mm | 19.7525 | 19.7134 | mm | 400 | 400 | 400 | -1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | 289.59 | 19.7809 | 0.067488 | -14.98 | 4 |  |  | - | - | - |  |  | - | NONE | 0.1 | 40 | 0 | 0.1 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020773; motor marked W002048/03; | Step_per_unit:400;Offset:289,59;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 285271100 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 30mm |  | mm | 0
3 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | a_m4_baff_hl | Y | N | - | W032652 | W032653 |  | - |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.6 | 2.4 | 0.4 | 4.8 | 2.512 | 24 | 12 | 1.5 | 50 |  | N | N | INV | N | - | N/A | absolute | 1493769300 | 0 |  | N/A | 200 | 1000000 | 200 | 2 | -5434 | 1488648723 | 1864 | 1503859033 | 1488770800 | 1503773200 | 1488648723 | 1503859033 |  | 400 | 0.15 | 18.245 | 1000000 | 833670.0466 | 1000000 | 20mm | 18.245 | 15.2103 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -1488.73 | -0.081277 | 15.129 | -5 | 10 |  |  | - | - | - |  |  | - | NONE | 0.1 | 40 | 0 | 0.1 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020773; motor marked W002035/03; | Step_per_unit:400;Offset:-1488,73;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 1493769300 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 20mm |  | mm | 0
4 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | a_m4_baff_hr | Y | N | - | W032654 | W032655 |  | - |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.6 | 2.4 | 0.4 | 4.8 | 2.512 | 24 | 12 | 1.5 | 50 |  | N | N | INV | N | - | N/A | absolute | 1731444100 | 0 |  | N/A | 200 | 1000000 | 200 | 2 | -4055 | 1726168125 | 2273 | 1741991012 | 1726440100 | 1741418400 | 1726168125 | 1741991012 |  | 400 | 0.15 | 15.82 | 1000000 | 1000182.4905 | 1000000 | 20mm | 15.82 | 15.8229 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -1726.09 | 0.078125 | 15.901 | -5 | 9.97 |  |  | - | - | - |  |  | - | NONE | 0.1 | 40 | 0 | 0.1 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020773; motor marked W002047/03; | Step_per_unit:400;Offset:-1726,09;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 1731444100 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 20mm |  | mm | 0
5 | B316A-EA03-DIA-BPM-01 | M4 BPM VERITAS | V | a_m4_bpm_v | Y | N | 2O | W032656 | W032657 |  | - |  | Oriental PKP268MD28B-L | Renishaw Resolute 32 bits BISS-C | 6 | 1.23 | 5.4 | 1.6 | 0.37 | 6.24 | 16.5321 | 48 | 12 | 1.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | - | 0 |  | N/A | 400 | 1000000 | 400 | 2 | -5331 | 23793762 | 101616 | 156798879 | - | - | 23793762 | 156798879 |  | 1600 | 0.3 | 66.8419 | 1000000 | 994923.5939 | 1000000 | 50mm | 133.6838 | 133.0051 | mm | 800 | 800 | 800 | -1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | 0 | -23.7938 | -156.7989 | - | - |  |  | - | - | - |  |  | - | DISABLE | 0.1 | 40 | 0 | 0.05 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020773; | Step_per_unit:800;Offset:0;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | - | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 50mm |  | mm | 0
6 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | a_m4_v1 | Y | N | 2O | W032658 | W032659 |  | - |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 7 | 3 | -4.5 | 2.4 | 4.396 | 24 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84229529 | 0 |  | N/A | 40 | 1000 | 400 | 2 | -405570 | 79160740 | 423303 | 89529602 | 79262957 | 89556843 | 79160740 | 89529602 |  | 1600 | 0.3 | 518.0456 | 10000 | 10007.6726 | 10000 | 10mm (±5mm) | 10360.9125 | 10368.862 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84229.529 | -5068.789 | 5300.073 | 79.263 | 89.5568 |  |  | - | - | - |  |  | - | DISABLE | 0.1 | 40 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020771; motor marked 11; | Step_per_unit:80;Offset:-84229,529;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | 0 |  | 84229529 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 10mm (±5mm) |  | um | 0
7 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | a_m4_v2 | Y | N | 2O | W032660 | W032661 |  | - |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 7.1 | 2.8 | -4.7 | 2.24 | 4.4588 | 24 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84265053 | 0 |  | N/A | 40 | 1000 | 400 | 2 | -431338 | 78870502 | 451171 | 89899058 | 78917342 | 89182222 | 78870502 | 89899058 |  | 1600 | 0.3 | 551.5681 | 10000 | 9997.4559 | 10000 | 10mm (±5mm) | 11031.3625 | 11028.556 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84265.053 | -5394.551 | 5634.005 | 78.9173 | 89.1822 |  |  | - | - | - |  |  | - | DISABLE | 0.1 | 40 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020771; motor marked 18; | Step_per_unit:80;Offset:-84265,053;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | 0 |  | 84265053 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 10mm (±5mm) |  | um | 0
8 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | a_m4_v3 | Y | N | 2O | W032662 | W032663 |  | - |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 7.2 | 2.9 | -4.6 | 2.32 | 4.5216 | 24 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84179513 | 0 |  | N/A | 40 | 1000 | 400 | 2 | -402520 | 79150746 | 411384 | 89323302 | 78778910 | 89084464 | 79150746 | 89323302 |  | 1600 | 0.3 | 508.69 | 10000 | 9998.7773 | 10000 | 10mm (±5mm) | 10173.8 | 10172.556 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84179.513 | -5028.767 | 5143.789 | 78.7789 | 89.0845 |  |  | - | - | - |  |  | - | DISABLE | 0.1 | 40 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020771; motor marked 35; | Step_per_unit:80;Offset:-84179,513;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | 0 |  | 84179513 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 10mm (±5mm) |  | um | 0
11 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | a_m4_h4 | Y | N | 2O | W032664 | W032665 |  | - |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 7.1 | 2.9 | -4.6 | 2.32 | 4.4588 | 24 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84516617 | 0 |  | N/A | 40 | 1000 | 400 | 2 | -404883 | 5409 | 415109 | 89699463 | 79454625 | 89489938 | 5409 | 89699463 |  | 1600 | 0.3 | 512.495 | 10000 | 87507.2479 | 10000 | 10mm (±5mm) | 10249.9 | 89694.054 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84516.617 | -84511.208 | 5182.846 | 79.4546 | 89.4899 |  |  | - | - | - |  |  | - | DISABLE | 0.1 | 40 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020771; motor marked 38; | Step_per_unit:80;Offset:-84516,617;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE
12 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | a_m4_h5 | Y | N | 2O | W032666 | W032667 |  | - |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 7 | 2.9 | -4.6 | 2.32 | 4.396 | 24 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84135330 | 0 |  | N/A | 40 | 1000 | 400 | 2 | -424114 | 78842287 | 435322 | 89574879 | 79187845 | 89575714 | 78842287 | 89574879 |  | 1600 | 0.3 | 537.1475 | 10000 | 9990.3583 | 10000 | 10mm (±5mm) | 10742.95 | 10732.592 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84135.33 | -5293.043 | 5439.549 | 79.1878 | 89.5757 |  |  | - | - | - |  |  | - | DISABLE | 0.1 | 40 | 0 | 0.5 | 0 | 0 | 0 | 0 | 0.1 | 50 | 0 |  |  | S020771; motor marked 41; | Step_per_unit:80;Offset:-84135,33;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE
13 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | X | a_mp1_x | Y | N | OPEN | W050216 | W050217 |  | #N/A |  | Oriental PK245M-02B series | Renishaw Resolute 32 bits BISS-C | 39.6 | 15 | 37 | 15.2 | 0.2 | 12.16 | 23.236 | 48 | 12 | 0.8 | 50 |  | N | N | INV | N | OPEN | N/A | absolute | 0 | 0 |  | N/A | 200 | 1000000 | 400 | 2 | 25409 | 63546787 | 37468 | 93644801 | 0 | 0 | 63546787 | 93644801 |  | 100 | 0.5 | 120.59 | 2000000 | 1996717.0744 | 2000000 | ±15mm | 30.1475 | 30.098 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE |  | 63.5468 | 93.6448 | 0 | 0 | 63.5468 | 93.6448 | 0 | 0 | [19 empty] | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
14 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Z | a_mp1_z | Y | N | OPEN | W050218 | W050219 |  | #N/A |  | Oriental PK245M-02B series | Renishaw Resolute 32 bits BISS-C | 39.6 | 15 | 38.66 | 15.5 | 0.5 | 12.4 | 24.2785 | 48 | 12 | 0.8 | 50 |  | N | N | INV | N | OPEN | N/A | absolute | 0 | 0 |  | N/A | 200 | 1000000 | 400 | 2 | 2591 | 6504094 | 14536 | 36322435 | 0 | 0 | 6504094 | 36322435 |  | 100 | 0.5 | 119.45 | 2000000 | 1997042.5115 | 2000000 | ±15mm | 29.8625 | 29.8183 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE |  | 6.5041 | 36.3224 | 0 | 0 | 6.5041 | 36.3224 | 0 | 0 | [19 empty] | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
15 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Y | a_mp1_y | Y | N | OPEN | W050220 | W050221 |  | #N/A |  | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.249 | 1 | 0.3 | 3.9 | 3.8238 | 48 | 12 | 3.5 | 90 |  | N | N | INV | N | OPEN | N/A | absolute | 0 | 0 |  | N/A | 200 | 1000000 | 400 | 2 | -27611 | 58251519 | 209304 | 419605289 | 0 | 0 | 58251519 | 419605289 |  | 400 | 2 | 592.2875 | 2000000 | 1220197.1846 | 2000000 | 358mm(400mm) | 592.2875 | 361.3538 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE |  | 58.2515 | 419.6053 | 0 | 0 | 58.2515 | 419.6053 | 0 | 0 | [19 empty] | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
16 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | YAW | a_mp1_yaw | Y | N | OPEN | W050222 | W050223 |  | #N/A |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 3.096 | 2.2 | 0.2 | 4.4 | 4.8607 | 48 | 12 | 1.6 | 70 |  |  |  | INV |  | OPEN | N/A | absolute | 0 | 0 |  | N/A | 200 | 11930445.5027 | 200 | 2 | -134242 | -938480000 | 171136 | 1183926000 | 0 | 0 | -938480000 | 1183926000 |  | 2000 | 0.5 | 152.689 | 11930445.5027 | 2780037.8547 | 2800000 | 0 | 763.445 | 177.8983 | deg | 400 | 1704.3494 | 400 | 1 | attr://EncAbsEnc | VALUE/11930445,5027191 | TRUE |  | -78.6626 | 99.2357 | 0 | 0 | -78.6626 | 99.2357 | 0 | 0 | [19 empty] | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/11930445,5027191;UseEncoderSource:TRUE
17 | B316A-EA05-DIA-SPC-01 | Q-Chamber VERITAS | YAW | a_spc1_yaw | Y | N | 0 | #N/A | #N/A |  | #N/A |  | Oriental PK296-E4.5 parallel | 0 | 1.5 | 0.24 |  |  | - | 0 | 0 | [9 empty] | 0 |  | absolute | 0 |  |  |  | 200 | 0 | 200 | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! | 0 |  |  | 0 |  |  | deg | 400 | 0 | [8 empty] | 0 | 0 | 0 | 0 | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
18 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VT | a_gr_baff_vt | Y | N | 0 | #N/A | #N/A |  | #N/A |  | Oriental PK245M-02B series | Renishaw Resolute 32 bits BISS-C | 39.6 | 15 |  |  | - | 0 | 0 | [9 empty] | 0 |  | absolute | 0 |  |  |  | 400 | 1000000 | 400 | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! | 1000000 |  |  | 0 |  |  | mm | 800 | 0 | [8 empty] | 0 | 0 | 0 | 0 | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
0
0
21 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VB | a_gr_baff_vb | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
22 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HR | a_gr_baff_hr | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
23 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HL | a_gr_baff_hl | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
24 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y1 | a_gra1_y1 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
25 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y2 | a_gra1_y2 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
26 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y3 | a_gra1_y3 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
27 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y1 | a_gra2_y1 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
28 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y2 | a_gra2_y2 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
0
0
31 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y3 | a_gra2_y3 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
32 | B316A-EA08-OPT-FLT-01 | Filter Unit VERITAS | X | a_flt1_x | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
33 | B316A-EA08-DIA-DETMC-01 | Detector VERITAS | Y | a_detmc1_y | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
34 | B316A-EA08-DIA-DETMC-02 | Detector VERITAS | Y | a_detmc2_y | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
35 | B316A-EA09-OPT-MM-01 | Movable Mask VERITAS | Y | a_mm1_y | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
36 | B316A-EA09-OPT-MM-02 | Movable Mask VERITAS | Y | a_mm2_y | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
37 | B316A-EA09-OPT-POL-01 | Multilayer Mirror VERITAS | Y | a_pol1_y | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
38 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | PIT | a_pol2_pit | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | deg | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
41 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | X | a_pol2_x | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
42 | B316A-EA09-DIA-TABLED-01 | Detector Table VERITAS | Z | a_tabled_z | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
43 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | PIT | a10_detmc1_pit | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | deg | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
44 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | Y | a10_detmc1_y | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | #N/A | #N/A | [29 empty] | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 |  |  | 1000000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
45 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT1 | a_detmc2_pit1 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 39.6 | 15 | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 1000000 | 0 |  |  | deg | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
46 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT2 | a_detmc2_pit2 | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 1.7 | 0.7 | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | deg | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
47 | B316A-EA11-DIA-SPC-01 | Spectrometer VERITAS | YAW | a11_spc1_yaw | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 1.5 | 0.24 | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | deg | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
48 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | X | a_mp2_x | Y | N | 0 | #N/A | #N/A |  | #N/A |  |  |  | 19.2 | 4.5 | [29 empty] | 0 | 0 | 0 | 0 |  | 1600 | 0.3 | 0 |  |  | 10000 | 0 |  |  | mm | 0 | 0 | [33 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
-13288610
-13288610
51 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Z | a_mp2_z | Y | N | 0 | #N/A | #N/A |  | #N/A |  | Oriental PK245M-02B series | 0 | 39.6 | 15 |  |  | - | 0 | 0 | [9 empty] | 0 |  | 0 | 0 | 0 | 0 | N/A | 400 | 0 | 400 | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 | 0 |  | 1000000 | 0 |  |  | mm | 800 | 0 | [6 empty] | #DIV/0! | #DIV/0! | 0 | 0 |  |  | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
52 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Y | a_mp2_y | Y | N | 0 | #N/A | #N/A |  | #N/A |  | Oriental PK264M-E2.0 parallel | 0 | 1.7 | 0.7 |  |  | - | 0 | 0 | [9 empty] | 0 |  | 0 | 0 | 0 | 0 | N/A | 400 | 0 | 400 | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  | 400 | 0.15 | 0 | 0 |  | 1000000 | 0 |  |  | mm | 800 | 0 | [6 empty] | #DIV/0! | #DIV/0! | 0 | 0 |  |  | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
53 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | b_m4_baff_vt | Y | N | - | W050829 | W050830 |  | #N/A |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.7 | 2.3 | 0.3 | 4.6 | 2.669 | 48 | 12 | 1.4 | 50 |  | N | N | INV | N | - | N/A | absolute | 0 | 0 | 0 | N/A | 200 | 1000000 | 200 | 2 | 224967 | 562462456 | 233141 | 582853230 | 0 | 0 | 562462456 | 582853230 |  | 200 | 0.3 | 40.87 | 1000000 | 997835.772 | 1000000 | 30mm | 20.435 | 20.3908 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -566.46 | -3.9975 | 16.3932 | 0 | 0 | -3.9975 | 16.3932 | - | - |  |  |  |  | NONE | [13 empty] | S035537; 692-23-50-00; 2018-04-20 SAT with movable rack | Step_per_unit:400;Offset:-566,46;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
1 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | b_m4_baff_vb | Y | N | - | W050831 | W050832 |  | #N/A |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.7 | 2.3 | 0.3 | 4.6 | 2.669 | 48 | 12 | 1.4 | 50 |  | N | N | INV | N | - | N/A | absolute | 0 | 0 | 0 | N/A | 200 | 1000000 | 200 | 2 | 420499 | 1051146688 | 428699 | 1071650857 | 0 | 0 | 1051146688 | 1071650857 |  | 200 | 0.3 | 41 | 1000000 | 1000203.3659 | 1000000 | 30mm | 20.5 | 20.5042 | mm | 400 | 400 | 400 | -1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | 1067.65 | 16.5033 | -4.0009 | 0 | 0 | 16.5033 | -4.0009 | - | - |  |  |  |  | NONE | [13 empty] | S035537; 692-23-50-00; 2018-04-20 SAT with movable rack | Step_per_unit:400;Offset:1067,65;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
55 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | b_m4_baff_hl | Y | N | - | W050833 | W050834 |  | #N/A |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.7 | 2.2 | 0.2 | 4.4 | 2.669 | 48 | 12 | 1.4 | 50 |  | N | N | INV | N | - | N/A | absolute | 0 | 0 | 0 | N/A | 200 | 1000000 | 200 | 2 | 297090 | 742731260 | 303872 | 759685446 | 0 | 0 | 742731260 | 759685446 |  | 200 | 0.3 | 33.91 | 1000000 | 999951.9906 | 1000000 | 20mm | 16.955 | 16.9542 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -747.73 | -4.9987 | 11.9554 | 0 | 0 | -4.9987 | 11.9554 | - | - |  |  |  |  | NONE | [13 empty] | S035537; 692-23-50-00; 2018-04-20 SAT with movable rack | Step_per_unit:400;Offset:-747,73;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
56 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | b_m4_baff_hr | Y | N | - | W050835 | W050836 |  | #N/A |  | Nanotec L4118S1404-M6X1 | Renishaw Resolute 32 bits BISS-C | 2.8 | 2 | 1.709 | 2.2 | 0.2 | 4.4 | 2.6831 | 48 | 12 | 1.4 | 50 |  | N | N | INV | N | - | N/A | absolute | 0 | 0 | 0 | N/A | 200 | 1000000 | 200 | 2 | 273691 | 684242398 | 280370 | 700940813 | 0 | 0 | 684242398 | 700940813 |  | 200 | 0.3 | 33.395 | 1000000 | 1000054.7986 | 1000000 | 20mm | 16.6975 | 16.6984 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -689.24 | -4.9976 | 11.7008 | 0 | 0 | -4.9976 | 11.7008 | - | - |  |  |  |  | NONE | [13 empty] | S035537; 692-23-50-00; 2018-04-20 SAT with movable rack | Step_per_unit:400;Offset:-689,24;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
57 | B316A-EB03-DIA-BPM-01 | M4 BPM VERITAS | V | b_m4_bpm_v | Y | N | 2O | W050837 | W050838 |  | #N/A |  | Oriental PKP268MD28B-L | Renishaw Resolute 32 bits BISS-C | 6 | 1.23 | 5.4 | 1.6 | 0.37 | 1.968 | 5.214 | 48 | 12 | 1.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 0 | 0 | 0 | N/A | 400 | 1000000 | 400 | 2 | 25851 | 32271641 | 126545 | 158192263 | 0 | 0 | 32271641 | 158192263 |  | 800 | 0.3 | 125.8675 | 1000000 | 1000422.047 | 1000000 | 50mm | 125.8675 | 125.9206 | mm | 800 | 800 | 800 | -1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | 158.19 | 125.9184 | -0.002263 | 0 | 0 | 125.9184 | -0.002263 | - | - |  |  |  |  | DISABLE | [13 empty] | S035537; 692-23-50-00; 2018-04-20 SAT with movable rack | Step_per_unit:800;Offset:158,19;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE
58 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | b_m4_v1 | Y | N | 2O | W050839 | W050840 |  | #N/A |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 6.83 | 3 | -4.5 | 22.5 | 40.2116 | 48 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84095873 | 0 | 0 | N/A | 40 | 1000 | 400 | 2 | 66956916 | 79194516 | 67841065 | 90254839 | 79056547 | 89155784 | 79194516 | 90254839 |  | 1600 | 0.8 | 552.5931 | 10000 | 10007.6553 | 10000 | 10mm (±5mm) | 11051.8625 | 11060.323 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84095.873 | -4901.357 | 6158.966 | 79056.547 | 89155.784 | -4901.357 | 6158.966 | - | - |  |  |  |  | DISABLE | [13 empty] | S035536; 692-23-06-00; mark 251; 2018-04-20 SAT with movable rack | Step_per_unit:80;Offset:-84095,873;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE
-13288610
-13288610
61 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | b_m4_v2 | Y | N | 2O | W050841 | W050842 |  | #N/A |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 6.873 | 3.1 | -4.4 | 23.25 | 40.4648 | 48 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 78960638 | 0 | 0 | N/A | 40 | 1000 | 400 | 2 | 66892060 | 79107633 | 67766951 | 90045882 | 73779002 | 83968939 | 79107633 | 90045882 |  | 1600 | 0.8 | 546.8069 | 10000 | 10001.9308 | 10000 | 10mm (±5mm) | 10936.1375 | 10938.249 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84209 | -5101.367 | 5836.882 | 73779.002 | 83968.939 | -5101.367 | 5836.882 | - | - |  |  |  |  | DISABLE | [13 empty] | S035536; 692-23-06-00; mark 248; 2018-04-20 SAT with movable rack | Step_per_unit:80;Offset:-84209;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE
62 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | b_m4_v3 | Y | N | 2O | W050843 | W050844 |  | #N/A |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 6.857 | 3 | -4.5 | 22.5 | 40.3706 | 48 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84128024 | 0 | 0 | N/A | 40 | 1000 | 400 | 2 | 66880768 | 79100318 | 67687061 | 89167257 | 79004220 | 89190308 | 79100318 | 89167257 |  | 1600 | 0.8 | 503.9331 | 10000 | 9988.368 | 10000 | 10mm (±5mm) | 10078.6625 | 10066.939 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84128.024 | -5027.706 | 5039.233 | 79004.22 | 89190.308 | -5027.706 | 5039.233 | - | - |  |  |  |  | DISABLE | [13 empty] | S035536; 692-23-06-00; mark 247; 2018-04-20 SAT with movable rack | Step_per_unit:80;Offset:-84128,024;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE
63 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | b_m4_h4 | Y | N | 2O | W050845 | W050846 |  | #N/A |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 6.9 | 3 | -4.5 | 22.5 | 40.6238 | 48 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 84001007 | 0 | 0 | N/A | 40 | 1000 | 400 | 2 | 66892052 | 78792192 | 67757608 | 89621535 | 78970165 | 89093371 | 78792192 | 89621535 |  | 1600 | 0.8 | 540.9725 | 10000 | 10009.1437 | 10000 | 10mm (±5mm) | 10819.45 | 10829.343 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84001.007 | -5208.815 | 5620.528 | 78970.165 | 89093.371 | -5208.815 | 5620.528 | - | - |  |  |  |  | DISABLE | [13 empty] | S035536; 692-23-06-00; mark 249; 2018-04-20 SAT with movable rack | Step_per_unit:80;Offset:-84001,007;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE
64 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | b_m4_h5 | Y | N | 2O | W050847 | W050848 |  | #N/A |  | Oriental PK245M-02B unipolar 1 coil | Renishaw Resolute 32 bits BISS-C | 9.9 | 7.5 | 6.9 | 3 | -4.5 | 22.5 | 40.6238 | 48 | 12 | 0.8 | 50 |  | N | N | INV | N | 2O | OT- OK, OT+ OK | absolute | 81975505 | 0 | 0 | N/A | 40 | 1000 | 400 | 2 | 66892882 | 79112207 | 67712667 | 89362478 | 76941717 | 87075481 | 79112207 | 89362478 |  | 1600 | 0.8 | 512.3656 | 10000 | 10002.8871 | 10000 | 10mm (±5mm) | 10247.3125 | 10250.271 | um | 80 | 80 | 80 | 1 | attr://EncAbsEnc | VALUE/1000 | TRUE | -84175.9 | -5063.693 | 5186.578 | 76941.717 | 87075.481 | -5063.693 | 5186.578 | - | - |  |  |  |  | DISABLE | [13 empty] | S035536; 692-23-06-00; mark 250; 2018-04-20 SAT with movable rack | Step_per_unit:80;Offset:-84175,9;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE
65 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A |  | #N/A |  | 0 | 0 | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | [9 empty] | 0 |  | 0 | 0 |  |  |  | 0 | 0 | #N/A | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! |  |  |  | 0 |  |  | mm | 0 | 0 | [6 empty] | #DIV/0! | #DIV/0! | 0 | 0 |  |  | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
66 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A |  | #N/A |  | 0 | 0 | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | [9 empty] | 0 |  | 0 | 0 |  |  |  | 0 | 0 | #N/A | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! |  |  |  | 0 |  |  | um | 0 | 0 | [6 empty] | #DIV/0! | #DIV/0! | 0 | 0 |  |  | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
67 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A |  | #N/A |  | 0 | 0 | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | [9 empty] | 0 |  | 0 | 0 |  |  |  | 0 | 0 | #N/A | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! |  |  |  | 0 |  |  | um | 0 | 0 | [6 empty] | #DIV/0! | #DIV/0! | 0 | 0 |  |  | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
68 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | #N/A | #N/A |  | #N/A |  | 0 | 0 | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | [9 empty] | 0 |  | 0 | 0 |  |  |  | 0 | 0 | #N/A | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! |  |  |  | 0 |  |  | um | 0 | 0 | [6 empty] | #DIV/0! | #DIV/0! | 0 | 0 |  |  | 0 | 0 | [19 empty] | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:
-13288610
-13288610
Note on AK1: encoder value after SAT

## Tango
axis | Equipment code | Equiment name | Axis name | Controller dns name | controller | Instrument | devicename | name | Steps per unit | attributes | unit | Soft lim- (units) | Soft lim+ (units) | Ready. (N or blank: don't take the data here)
1 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/baff-01-vt | a_m4_baff_vt | 400 | Step_per_unit:400;Offset:-1137,23;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 0 | 0
2 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/baff-01-vb | a_m4_baff_vb | 400 | Step_per_unit:400;Offset:289,59;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 0 | 0
3 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/baff-01-hl | a_m4_baff_hl | 400 | Step_per_unit:400;Offset:-1488,73;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 0 | 0
4 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/baff-01-hr | a_m4_baff_hr | 400 | Step_per_unit:400;Offset:-1726,09;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 0 | 0
5 | B316A-EA03-DIA-BPM-01 | M4 BPM VERITAS | V | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/dia/bpm-01-v | a_m4_bpm_v | 800 | Step_per_unit:800;Offset:0;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 0 | 0
6 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/mir-04-v1 | a_m4_v1 | 80 | Step_per_unit:80;Offset:-84229,529;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | 0 | 0
7 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/mir-04-v2 | a_m4_v2 | 80 | Step_per_unit:80;Offset:-84265,053;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | 0 | 0
8 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/mir-04-v3 | a_m4_v3 | 80 | Step_per_unit:80;Offset:-84179,513;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | 0 | 0
0 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0
11 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/mir-04-h4 | a_m4_h4 | 80 | Step_per_unit:80;Offset:-84516,617;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | 0 | 0
12 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA03 | b316a-ea03/opt/mir-04-h5 | a_m4_h5 | 80 | Step_per_unit:80;Offset:-84135,33;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | 0 | 0
13 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | X | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA04 | b316a-ea04/dia/mp-01-x | a_mp1_x | 400 | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 63.5468 | 93.6448
14 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Z | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA04 | b316a-ea04/dia/mp-01-z | a_mp1_z | 400 | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 6.5041 | 36.3224
15 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA04 | b316a-ea04/dia/mp-01-y | a_mp1_y | 400 | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | 58.2515 | 419.6053
16 | B316A-EA04-DIA-MP-01 | Manipulator VERITAS | YAW | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA04 | b316a-ea04/dia/mp-01-yaw | a_mp1_yaw | 400 | Step_per_unit:400;Offset:;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/11930445,5027191;UseEncoderSource:TRUE | deg | -78.6626 | 99.2357
17 | B316A-EA05-DIA-SPC-01 | Q-Chamber VERITAS | YAW | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA05 | b316a-ea05/dia/spc-01-yaw | a_spc1_yaw | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | deg | 0 | 0
18 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | VT | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/baff-01-vt | a_gr_baff_vt | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
0
0
21 | § | Grating Baffles VERITAS | VB | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 |  | a_gr_baff_vb | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
22 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HR | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/baff-01-hr | a_gr_baff_hr | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
23 | B316A-EA07-OPT-BAFF-01 | Grating Baffles VERITAS | HL | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/baff-01-hl | a_gr_baff_hl | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
24 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y1 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/gra-01-y1 | a_gra1_y1 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
25 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y2 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/gra-01-y2 | a_gra1_y2 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
26 | B316A-EA07-OPT-GRA-01 | Grating VERITAS | Y3 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/gra-01-y3 | a_gra1_y3 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
27 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y1 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/gra-02-y1 | a_gra2_y1 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
28 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y2 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/gra-02-y2 | a_gra2_y2 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
0
0
31 | B316A-EA07-OPT-GRA-02 | Grating VERITAS | Y3 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA07 | b316a-ea07/opt/gra-02-y3 | a_gra2_y3 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
32 | B316A-EA08-OPT-FLT-01 | Filter Unit VERITAS | X | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA08 | b316a-ea08/opt/flt-01-x | a_flt1_x | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
33 | B316A-EA08-DIA-DETMC-01 | Detector VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA08 | b316a-ea08/dia/detmc-01-y | a_detmc1_y | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
34 | B316A-EA08-DIA-DETMC-02 | Detector VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA08 | b316a-ea08/dia/detmc-02-y | a_detmc2_y | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
35 | B316A-EA09-OPT-MM-01 | Movable Mask VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA09 | b316a-ea09/opt/mm-01-y | a_mm1_y | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
36 | B316A-EA09-OPT-MM-02 | Movable Mask VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA09 | b316a-ea09/opt/mm-02-y | a_mm2_y | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
37 | B316A-EA09-OPT-POL-01 | Multilayer Mirror VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA09 | b316a-ea09/opt/pol-01-y | a_pol1_y | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
38 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | PIT | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA09 | b316a-ea09/opt/pol-02-pit | a_pol2_pit | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | deg | 0 | 0
41 | B316A-EA09-OPT-POL-02 | Multilayer Mirror VERITAS | X | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA09 | b316a-ea09/opt/pol-02-x | a_pol2_x | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
42 | B316A-EA09-DIA-TABLED-01 | Detector Table VERITAS | Z | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA09 | b316a-ea09/dia/tabled-01-z | a_tabled_z | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
43 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | PIT | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA10 | b316a-ea10/dia/detmc-01-pit | a10_detmc1_pit | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | deg | 0 | 0
44 | B316A-EA10-DIA-DETMC-01 | Detector VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA10 | b316a-ea10/dia/detmc-01-y | a10_detmc1_y | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
45 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT1 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA10 | b316a-ea10/dia/detmc-02-pit1 | a_detmc2_pit1 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | deg | 0 | 0
46 | B316A-EA10-DIA-DETMC-02 | Detector VERITAS | PIT2 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA10 | b316a-ea10/dia/detmc-02-pit2 | a_detmc2_pit2 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | deg | 0 | 0
47 | B316A-EA11-DIA-SPC-01 | Spectrometer VERITAS | YAW | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA11 | b316a-ea11/dia/spc-01-yaw | a11_spc1_yaw | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | deg | 0 | 0
48 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | X | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA04 | b316a-ea04/dia/mp-02-x | a_mp2_x | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
51 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Z | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA04 | b316a-ea04/dia/mp-02-z | a_mp2_z | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
52 | B316A-EA04-DIA-MP-02 | Manipulator VERITAS | Y | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EA04 | b316a-ea04/dia/mp-02-y | a_mp2_y | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
53 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VT | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/baff-01-vt | b_m4_baff_vt | 400 | Step_per_unit:400;Offset:-566,46;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -3.9975 | 16.3932
54 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | VB | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/baff-01-vb | b_m4_baff_vb | 400 | Step_per_unit:400;Offset:1067,65;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -4.0009 | 16.5033
55 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HL | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/baff-01-hl | b_m4_baff_hl | 400 | Step_per_unit:400;Offset:-747,73;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -4.9987 | 11.9554
56 | B316A-EB03-OPT-BAFF-01 | M4 Baffles VERITAS | HR | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/baff-01-hr | b_m4_baff_hr | 400 | Step_per_unit:400;Offset:-689,24;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -4.9976 | 11.7008
57 | B316A-EB03-DIA-BPM-01 | M4 BPM VERITAS | V | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/dia/bpm-01-y | b_m4_bpm_v | 800 | Step_per_unit:800;Offset:158,19;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -0.002263 | 125.9184
58 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/mir-04-v1 | b_m4_v1 | 80 | Step_per_unit:80;Offset:-84095,873;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | -4901.357 | 6158.966
61 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/mir-04-v2 | b_m4_v2 | 80 | Step_per_unit:80;Offset:-84209;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | -5101.367 | 5836.882
62 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/mir-04-v3 | b_m4_v3 | 80 | Step_per_unit:80;Offset:-84128,024;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | -5027.706 | 5039.233
63 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/mir-04-h4 | b_m4_h4 | 80 | Step_per_unit:80;Offset:-84001,007;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | -5208.815 | 5620.528
64 | B316A-EB03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5 | b316a-a101630-cab05-ctl-ipap-01 | icepap_ctrl_2 | /EB03 | b316a-eb03/opt/mir-04-h5 | b_m4_h5 | 80 | Step_per_unit:80;Offset:-84175,9;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000;UseEncoderSource:TRUE | um | -5063.693 | 5186.578
65 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm | 0 | 0
66 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | um | 0 | 0
67 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | um | 0 | 0
68 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | Step_per_unit:;Offset:;Sign:;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | um | 0 | 0

## MOTION1
Cable ID | Conf. Code | Equip. A | Type A | Chan. A | Equip. B | Type B | Chan. B | Chan. B cleaned | AUX
W032648 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML1 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VTML | VT | B316A-EA03-OPT-BAFF-01VT
W032650 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML2 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VBML | VB | B316A-EA03-OPT-BAFF-01VB
W032652 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML3 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HLML | HL | B316A-EA03-OPT-BAFF-01HL
W032654 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML4 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HRML | HR | B316A-EA03-OPT-BAFF-01HR
W032656 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML5 | B316A-EA03-DIA-BPM-01 | BPM Veritas | VML | V | B316A-EA03-DIA-BPM-01V
W032658 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML6 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V1ML | V1 | B316A-EA03-OPT-MIR-04V1
W032660 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML7 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V2ML | V2 | B316A-EA03-OPT-MIR-04V2
W032662 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML8 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V3ML | V3 | B316A-EA03-OPT-MIR-04V3
W032664 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML1 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H4ML | H4 | B316A-EA03-OPT-MIR-04H4
W032666 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML2 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H5ML | H5 | B316A-EA03-OPT-MIR-04H5
- | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML3 |  |  |  | 0 | 0
- | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML4 |  |  |  | 0 | 0
- | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML5 |  |  |  | 0 | 0
- | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML6 |  |  |  | 0 | 0
- | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML7 |  |  |  | 0 | 0
- | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML8 |  |  |  | 0 | 0
W032648 | 8 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML1 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VTML | VT | B316A-EA03-OPT-BAFF-01VT
W032650 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML2 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VBML | VB | B316A-EA03-OPT-BAFF-01VB
W032652 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML3 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HLML | HL | B316A-EA03-OPT-BAFF-01HL
W032654 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML4 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HRML | HR | B316A-EA03-OPT-BAFF-01HR
W032656 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML5 | B316A-EA03-DIA-BPM-01 | BPM Veritas | VML | V | B316A-EA03-DIA-BPM-01V
W032658 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML6 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V1ML | V1 | B316A-EA03-OPT-MIR-04V1
W032660 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML7 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V2ML | V2 | B316A-EA03-OPT-MIR-04V2
W032662 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | ML8 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V3ML | V3 | B316A-EA03-OPT-MIR-04V3
W032664 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML1 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H4ML | H4 | B316A-EA03-OPT-MIR-04H4
W032666 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML2 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H5ML | H5 | B316A-EA03-OPT-MIR-04H5
W050216 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML3 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | XML | X | B316A-EA04-DIA-MP-01X
W050218 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML4 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | ZML | Z | B316A-EA04-DIA-MP-01Z
W050220 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML5 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | YML | Y | B316A-EA04-DIA-MP-01Y
W050222 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML6 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | YAWML | YAW | B316A-EA04-DIA-MP-01YAW
W050829 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML3 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | VTML | VT | B316A-EB03-OPT-BAFF-01VT
W050831 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML4 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | VBML | VB | B316A-EB03-OPT-BAFF-01VB
W050833 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML5 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | HLML | HL | B316A-EB03-OPT-BAFF-01HL
W050835 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML6 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | HRML | HR | B316A-EB03-OPT-BAFF-01HR
W050837 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML7 | B316A-EB03-DIA-BPM-01 | BPM Veritas | VML | V | B316A-EB03-DIA-BPM-01V
W050839 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML8 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | V1ML | V1 | B316A-EB03-OPT-MIR-04V1
W050841 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML1 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | V2ML | V2 | B316A-EB03-OPT-MIR-04V2
W050843 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML2 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | V3ML | V3 | B316A-EB03-OPT-MIR-04V3
W050845 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML3 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | H4ML | H4 | B316A-EB03-OPT-MIR-04H4
W050847 | MOTION-01 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML4 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | H5ML | H5 | B316A-EB03-OPT-MIR-04H5

## MOTION3
Cable ID | Conf. Code | Equip. A | Type A | Chan. A | Equip. B | Type B | Chan. B | Chan. B cleaned | AUX
W032649 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC1 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VTEC | VT | B316A-EA03-OPT-BAFF-01VT
W032651 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC2 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VBEC | VB | B316A-EA03-OPT-BAFF-01VB
W032653 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC3 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HLEC | HL | B316A-EA03-OPT-BAFF-01HL
W032655 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC4 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HREC | HR | B316A-EA03-OPT-BAFF-01HR
W032657 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC5 | B316A-EA03-DIA-BPM-01 | BPM Veritas | VEC | V | B316A-EA03-DIA-BPM-01V
W032659 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC6 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V1EC | V1 | B316A-EA03-OPT-MIR-04V1
W032661 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC7 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V2EC | V2 | B316A-EA03-OPT-MIR-04V2
W032663 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC8 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V3EC | V3 | B316A-EA03-OPT-MIR-04V3
W032665 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC1 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H4EC | H4 | B316A-EA03-OPT-MIR-04H4
W032667 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC2 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H5EC | H5 | B316A-EA03-OPT-MIR-04H5
- | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC3 |  |  |  | 0 | 0
- | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC4 |  |  |  | 0 | 0
- | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC5 |  |  |  | 0 | 0
- | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC6 |  |  |  | 0 | 0
- | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC7 |  |  |  | 0 | 0
- | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC8 |  |  |  | 0 | 0
W032649 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC1 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VTEC | VT | B316A-EA03-OPT-BAFF-01VT
W032651 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC2 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | VBEC | VB | B316A-EA03-OPT-BAFF-01VB
W032653 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC3 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HLEC | HL | B316A-EA03-OPT-BAFF-01HL
W032655 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC4 | B316A-EA03-OPT-BAFF-01 | Baffle Veritas | HREC | HR | B316A-EA03-OPT-BAFF-01HR
W032657 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC5 | B316A-EA03-DIA-BPM-01 | BPM Veritas | VEC | V | B316A-EA03-DIA-BPM-01V
W032659 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC6 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V1EC | V1 | B316A-EA03-OPT-MIR-04V1
W032661 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC7 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V2EC | V2 | B316A-EA03-OPT-MIR-04V2
W032663 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP MASTER | EC8 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | V3EC | V3 | B316A-EA03-OPT-MIR-04V3
W032665 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC1 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H4EC | H4 | B316A-EA03-OPT-MIR-04H4
W032667 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC2 | B316A-EA03-OPT-MIR-04 | Mirror Veritas | H5EC | H5 | B316A-EA03-OPT-MIR-04H5
W050217 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC3 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | XEC | X | B316A-EA04-DIA-MP-01X
W050219 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC4 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | ZEC | Z | B316A-EA04-DIA-MP-01Z
W050221 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC5 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | YEC | Y | B316A-EA04-DIA-MP-01Y
W050223 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC6 | B316A-EA04-DIA-MP-01 | Manipulator Veritas | YAWEC | YAW | B316A-EA04-DIA-MP-01YAW
W050830 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC3 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | VTEC | VT | B316A-EB03-OPT-BAFF-01VT
W050832 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC4 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | VBEC | VB | B316A-EB03-OPT-BAFF-01VB
W050834 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC5 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | HLEC | HL | B316A-EB03-OPT-BAFF-01HL
W050836 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC6 | B316A-EB03-OPT-BAFF-01 | Baffle Veritas | HREC | HR | B316A-EB03-OPT-BAFF-01HR
W050838 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC7 | B316A-EB03-DIA-BPM-01 | BPM Veritas | VEC | V | B316A-EB03-DIA-BPM-01V
W050840 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC8 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | V1EC | V1 | B316A-EB03-OPT-MIR-04V1
W050842 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC1 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | V2EC | V2 | B316A-EB03-OPT-MIR-04V2
W050844 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC2 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | V3EC | V3 | B316A-EB03-OPT-MIR-04V3
W050846 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC3 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | H4EC | H4 | B316A-EB03-OPT-MIR-04H4
W050848 | MOTION-03 | B316A-A101630-CAB05-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC4 | B316A-EB03-OPT-MIR-04 | Mirror Veritas | H5EC | H5 | B316A-EB03-OPT-MIR-04H5
[8 empty] | 0 | 0
[8 empty] | 0 | 0
[8 empty] | 0 | 0

## DISABLE
Cable ID | Conf. Code | Equip. A | Type A | Chan. A | Equip. B | Type B | Chan. B | Chan. B cleaned | AUX
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS1 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VTDIS | VT | B316A-EA03-OPT-BAFF-01VT
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS2 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | VBDIS | VB | B316A-EA03-OPT-BAFF-01VB
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS3 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HLDIS | HL | B316A-EA03-OPT-BAFF-01HL
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS4 | B316A-EA03-OPT-BAFF-01 | M4 Baffles VERITAS | HRDIS | HR | B316A-EA03-OPT-BAFF-01HR
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS5 | B316A-EA03-DIA-BPM-01 | M4 BPM VERITAS | VDIS | V | B316A-EA03-DIA-BPM-01V
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS6 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V1DIS | V1 | B316A-EA03-OPT-MIR-04V1
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS7 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V2DIS | V2 | B316A-EA03-OPT-MIR-04V2
- | - | B316A-A101630-CAB05-CTL-IPAP-01 | ESRF IcePAP SLAVE | DIS8 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | V3DIS | V3 | B316A-EA03-OPT-MIR-04V3
- | - | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | DIS1 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H4DIS | H4 | B316A-EA03-OPT-MIR-04H4
- | - | B316A-A101630-CAB05-CTL-IPAP-02 | ESRF IcePAP SLAVE | DIS2 | B316A-EA03-OPT-MIR-04 | Focusing Mirror M4 VERITAS | H5DIS | H5 | B316A-EA03-OPT-MIR-04H5
DIS3
DIS4
DIS5
DIS6
DIS7
DIS8

## MOTORS
Number of motors: | 105
Motor model and wiring (if many) | Vnominal | Iphase(A) | Rphase(Ohm) | Lphase(mH) | full steps | NVOLT | IVOLT | NCURR | ICURR | MOTPOLES | ANTURN | ANSTEP | DEFVEL | DEFACCT | MREGMODE | CURRGAIN | HOMEVEL | ABSMODE | SSIDBITS | SSICODE | SSICLOCK | SSIDELAY | SSISTATUS | SSISTRTB | SSIMSKNB | SSIMSKMSB | SSILDC | SSIALDC | SSIEWCHK | SSIEWPOL | SSIEWPOS | SSIEECHK | SSIEEPOL | SSIEEPOS
Stoegra Sm 56.2.18 j3 | 60 | 5.9 | 1.2 | 2.6 | 200 | 60 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Stoegra SM 56.1.18.J1 | 60 | 2 | 4.75 | 9 | 200 | 60 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Stoegra SM 56.3.18.J4 | 55 | 7.5 | 0.72 | 1.2 | 200 | 55 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 19.200.0.6 parallel | 25 | 0.8 | 1.9 | 0.5 | 200 | 25 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 25.200.1.2 parallel | 25 | 1.7 | 0.95 | 0.4 | 200 | 25 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 32.200.1.2 parallel | 42 | 1.7 | 1.25 | 1.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 32.200.2.5 parallel | 25 | 3.5 | 0.3 | 0.3 | 200 | 25 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 41.500.0.6 parallel | 42 | 0.8 | 5.1 | 7.6 | 500 | 42 | 12 | 0.8 | 50 | 125 | 1 | 1000 | 2500 | 0.3 | CURR | LOW | 1250
Phytron ZSS 42.200.1.2 parallel | 42 | 1.7 | 1.7 | 3 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 43.500.2.5 parallel | 42 | 3.5 | 0.5 | 1.2 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 52.200.2.5 parallel | 48 | 3.5 | 0.6 | 1.6 | 200 | 48 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS.52.200.2.5 parallel | 48 | 3.5 | 0.6 | 1.6 | 200 | 48 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS.52.200.2.5 series | 48 | 1.8 | 2.4 | 6.4 | 200 | 48 | 12 | 1.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 57.200.1.2 parallel | 48 | 1.7 | 3.9 | 7.8 | 200 | 48 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 57.200.2.5 parallel | 48 | 3.5 | 0.8 | 2.4 | 200 | 48 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 57-3.200.6.5 parallel | 48 | 7.5 | 0.4 | 1.2 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 88/3.200.8 parallel | 55 | 7.5 | 0.14 | 2.6 | 200 | 55 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 87/2.200.6.5 parallel | 55 | 7.5 | 0.25 | 1.5 | 200 | 55 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 87/2.200.8.4 parallel | 55 | 7.5 | 0.15 | 1 | 200 | 55 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 88/2.200.8 parallel | 55 | 7.5 | 0.1 | 1.5 | 200 | 55 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 19.200.0.6 parallel | 25 | 0.8 | 1.05 | 0.55 | 200 | 25 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 25.200.1.2 parallel | 25 | 1.7 | 0.47 | 0.4 | 200 | 25 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 32.200.1.2 parallel | 42 | 1.7 | 0.65 | 1.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 33.200.1.2 parallel | 42 | 1.7 | 0.95 | 2.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 42.200.1.2 parallel | 42 | 1.7 | 0.85 | 3 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 42.500.2.5 parallel | 25 | 3.5 | 0.17 | 0.7 | 500 | 25 | 12 | 3.5 | 50 | 125 | 1 | 1000 | 2500 | 0.3 | CURR | LOW | 1250
Phytron VSS 43.200.1.2 parallel | 42 | 1.7 | 1.3 | 5.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 43.200.2.5 parallel | 42 | 3.5 | 0.25 | 1.2 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 52.200.1.2 parallel | 42 | 1.7 | 1.33 | 7 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 52.200.2.5 parallel | 42 | 3.5 | 0.3 | 1.6 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 52.200.5 parallel | 25 | 7 | 0.08 | 0.4 | 200 | 25 | 20 | 7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 57.200.1.2 parallel | 42 | 1.7 | 2 | 9.5 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 57.200.2.5 parallel | 42 | 3.5 | 0.4 | 2.4 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Superior Slo-syn KML062F13 | 48 | 7.5 | 0.2 | 0.85 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Superior Slo-syn KML061F05 | 48 | 3.8 | 0.85 | 4.56 | 200 | 48 | 12 | 3.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Superior Slo-syn KML061F02 | 48 | 1.5 | 4.94 | 30.1 | 200 | 48 | 12 | 1.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK223PB series | 48 | 0.9 | 5.6 | 4 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK243-01-SGxx series | 48 | 0.9 | 8.4 | 10 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK243-02-SGxx series | 48 | 0.4 | 48 | 60 | 200 | 48 | 12 | 0.4 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK243-01-SGxx series | 48 | 0.9 | 8.4 | 10 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK243-01-SGxx series | 48 | 0.9 | 8.4 | 10 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK244-01 series | 48 | 1.2 | 6.6 | 12.8 | 200 | 48 | 12 | 1.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK244-01 parallel | 48 | 2.4 | 1.8 | 3.2 | 200 | 48 | 12 | 2.4 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK245-01 series | 48 | 1.2 | 6.6 | 11.2 | 200 | 48 | 12 | 1.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK245DB | 48 | 2.1 | 2.1 | 3.1 | 200 | 48 | 12 | 2.1 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK245MD15B | 48 | 2.1 | 2.12 | 5.1 | 400 | 48 | 12 | 2.1 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK245M-02B series | 48 | 0.8 | 15 | 39.6 | 400 | 48 | 12 | 0.8 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK245M-02B unipolar 1 coil | 48 | 0.8 | 7.5 | 9.9 | 400 | 48 | 12 | 0.8 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK264-02 series | 48 | 2 | 2.8 | 5.6 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264-E2.0 series | 48 | 2 | 2.8 | 5.6 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264M-E2.0 parallel | 48 | 3.9 | 0.7 | 1.7 | 400 | 48 | 12 | 3.9 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK264BE-SGxx series | 48 | 2 | 2.8 | 5.6 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264BE-SGxx parallel | 48 | 3.9 | 0.7 | 1.4 | 200 | 48 | 12 | 3.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264JDA | 48 | 3.9 | 0.73 | 1.8 | 200 | 48 | 12 | 3.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK266-02 series | 48 | 2 | 3.6 | 10 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK266D14B | 48 | 2 | 3.6 | 10 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK267JA series | 48 | 2 | 4.8 | 14.2 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268-E2.0 series | 48 | 2 | 4.5 | 14.4 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268M-E2.0 parallel | 48 | 3.9 | 1.13 | 4.8 | 400 | 48 | 12 | 3.9 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK268-01 series | 48 | 1 | 17.2 | 56 | 200 | 48 | 12 | 1 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268-02 series | 48 | 2 | 4.5 | 14.4 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268D14B | 48 | 2 | 4.5 | 14.4 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268M-E2.0 series | 48 | 2 | 4.5 | 19.2 | 400 | 48 | 12 | 2 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK268M-02 series | 48 | 2 | 4.5 | 19.2 | 400 | 48 | 12 | 2 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK268-03 series | 48 | 2.9 | 2 | 6.4 | 200 | 48 | 12 | 2.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK269JD | 48 | 3.9 | 1.49 | 5.7 | 200 | 48 | 12 | 3.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296-E4.5 series | 48 | 4.5 | 0.96 | 6 | 200 | 48 | 12 | 4.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296-E4.5 parallel | 48 | 7.5 | 0.24 | 1.5 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296-AE-SGxx series | 48 | 2.9 | 0.96 | 6 | 200 | 48 | 12 | 2.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296-AE-SGxx parallel | 48 | 5.9 | 0.24 | 1.5 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK299-E4.5 series | 48 | 4.5 | 0.96 | 6 | 200 | 48 | 12 | 4.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK299-E4.5 parallel | 48 | 7.5 | 0.24 | 1.5 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PKP264MU20B-L unipolar 1 coil | 48 | 2 | 2 |  | 400 | 48 | 12 | 2 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PKP268MD28B-L | 48 | 3.9 | 1.23 | 6 | 400 | 48 | 12 | 3.9 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Faulhaber Precistep AM1524-2R-A-0,25-12,5 | 20 | 0.4 | 12.5 | 6.3 | 24 | 20 | 12 | 0.4 | 50 | 6 | 1 | 48 | 120 | 0.3 | CURR | LOW | 60
Faulhaber Precistep AM1524-2R-A-0,45-3,6 | 20 | 0.6 | 3.6 | 1.9 | 24 | 20 | 12 | 0.6 | 50 | 6 | 1 | 48 | 120 | 0.3 | CURR | LOW | 60
McLennan 23HT18C230 parallel | 48 | 5.9 | 0.38 | 1.2 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HS-309 series | 48 | 3.4 | 1.48 | 2.92 | 200 | 48 | 12 | 3.4 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HS-309 parallel | 48 | 6.6 | 0.37 | 0.73 | 200 | 48 | 20 | 6.6 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 17HS-240E | 48 | 3.2 | 0.72 | 0.83 | 200 | 48 | 12 | 3.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HS-108 Mk2 | 48 | 5.5 | 0.37 | 0.59 | 200 | 48 | 20 | 5.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HSX-206 parallel | 48 | 5.9 | 0.35 | 0.9 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HSX-306E parallel | 48 | 5.9 | 0.55 | 1.7 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Berger Lahr VRDM 268/50L3G8A | 48 | 4.3 | 2.2 | 20 | 200 | 48 | 12 | 4.3 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
PowerMaxII P21NRXD-LNF-NS-00 | 48 | 1.1 | 5 | 10 | 200 | 48 | 12 | 1.1 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear Engineering 3518X-04D-0 | 48 | 0.6 | 3.8 | 2.7 | 200 | 48 | 12 | 0.6 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear Engineering 5718X-18DE-01 series | 48 | 4.2 | 0.8 | 1.4 | 200 | 48 | 12 | 4.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear Engineering 5718X-18DE-01 parallel | 48 | 7.5 | 0.2 | 0.35 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear engineering 41185-04PD-23RO parallel | 48 | 1.9 | 2.5 | 3.1 | 200 | 48 | 12 | 1.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear engineering 4118S-04s series | 48 | 0.9 | 9.9 | 12.5 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear engineering 4118S-04s parallel | 48 | 1.9 | 2.5 | 3.1 | 200 | 48 | 12 | 1.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotec ST5918L3008 series | 48 | 2.9 | 2 | 8.8 | 200 | 48 | 12 | 2.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotec ST5918L3008 parallel | 48 | 5.9 | 0.5 | 2.2 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotec L4118S1404-M6X1 | 48 | 2 | 2 | 2.8 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Standa 8cma20-8/15 | 24 | 0.8 | 5.4 | 1.4 | 200 | 24 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Standa PG15/50 | 24 | 0.7 | 10 | 3 | 20 | 24 | 12 | 0.7 | 50 | 5 | 1 | 40 | 100 | 0.3 | CURR | LOW | 50
Fulling 28STH45B2-096 series | 24 | 0.8 | 6.8 | 4.9 | 200 | 24 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Trinamic Qmot QSH-4128-35-10-27 | 24 | 2 | 5.3 | 6.6 | 200 | 24 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Newport TraXXppd | 24 | 0.4 | 12.5 | 5.5 | 20 | 24 | 12 | 0.4 | 50 | 5 | 1 | 40 | 100 | 0.3 | CURR | LOW | 50
Renishaw Resolute 32 bits | [17 empty] | BISSC | 0 | BINARY | 1.25MHz | 0 | S | 0 | 1 | LSB | 0 | AUTO | YES | LOW | 6 | YES | LOW | 7
Renishaw Resolute 26 bits | [17 empty] | BISSC | 26 | BINARY | 1.25MHz | 0 | S | 0 | 0 | LSB | 0 | AUTO | YES | LOW | 6 | YES | LOW | 7
Renishaw Resolute 36 bits | [17 empty] | BISSC | 1 | BINARY | 1.25MHz | 0 | S | 0 | 5 | LSB | 0 | AUTO | YES | LOW | 6 | YES | LOW | 7
RLS LMA10DCxxxxxxxx | [17 empty] | BISSC | 26 | BINARY | 1.25MHz | 0 | S | 1 | 0 | LSB | 0 | AUTO | NO | LOW | 6 | YES | HIGH | 7