## Revision Log
Remove this sheet. It is just for in-document template purposes
Date | Who | Comments
150701 | Julio Lidón-Simón | Added Standa, Fulling, Trinamic, OM PK244-01 parallel motors. Added RLS LMA encoders
Additions to 'Tests' sheet:
Independent axis (Y/N)
Can axis collide (Y/N)
Motor phase mH
Motor phase Ohms
R Cable (Ohms)
Lim- position (ects or st) from FAT
Lim+ position (ects or st) from FAT
Time between limits (s)
Datasheet stroke (units)
Lim- position (ects prehoming) to compare with FAT
Lim+ position (ects prehoming) to compare with FAT
150706 | Julio Lidón-Simón | Columns of Motor R and L in the 'tests' sheet were inverted. Corrected
Added McLennan 23HS-309 series
Added 'DISABLE' cable sheet
Added Disable cable tests in the 'tests' sheet
150714 | Julio Lidón-Simón | In Sheet 'Tests' there's an extra column for the ects/motor turn configured (apart from the other two calculated from theoretical 'Data' and st/ects between limits
|  | Corrected: | Phytron VSS axes resistance
|  | Added: | Phytron VSS 25.200.0.6 parallel
Phytron VSS 33.200.0.6 parallel
McLennan 23HS-309 series
150911 | Julio Lidón-Simón | Motors sheet: Removed NVOLT and IVOLT
Motors sheet: Corrected Nanotech R and L values
Motors sheet: PKP268MD28x-x inductance updated
Added TANGO sheet
Sheet 'System info' changed to 'SystemInfo'.
Changes to new 'SystemInfo'. Added number of cables
Data sheet: Full change of structure
Added POSITIONS sheet
Added DISABLE sheet
Tests sheet: Renamed and added a big number of columns
List of steps to update to this version:
Update motor PKP268MD28x-x
+ Modify 'Data'
- Unhide hidden data columns (E) in 'Data' sheet
- Add two new columns in A and B shifting right everything.
- Copy in the empty lines at the end of your data in 'Data' the info from columns A-D in 'Tests'. The information in 'Data' will have that structure in the new layout.
- In the rest of the fields for those lines you just fill cols A-D fill the information from the lines already present in the 'Data' sheet (take the info from its equipment type).
- Rename the sheet 'System info' to 'SystemInfo'
- Once that is done, write the number of cables in SystemInfo!$B$1 from Data!$C$2 (initially was in A2 but was moved there when the two columns were added in step 2).
- Using Replace Within:Workbook, Look in: Formulas Find what: Data!$C$2, Replace with: SystemInfo!$B$1 update all the formulas so that Data!$C$1 is taken instead
- Remove all the rows in 'Data' between the first row and the row with the column titles (the column titles will be now in row 2).
- In the sheet 'Data' correct the formulas in the columns N-Q (taking info from 'MOTORS') so that they have NO spaces in the formulas :o (dirty one).
- Using Replace Within:Workbook, Look in: Formulas Find what: $1;0, Replace with: $2;0 update all the formulas (we're going to remove the first row from 'Data' too).
- Using Replace Within:Workbook, Look in: Formulas Find what:'MATCH(CONCATENATE($C' , Replace with: 'MATCH(CONCATENATE($B'
- At the end, Add column F (sardana name) and fill it with the scientist and column AE (TC tango device)
+ Create 'POSITIONS' sheet
- Fill columns till E
- Get columns F-G from the formulas
- Fill Columns I-N from the info in Data (just copy)
- Fill Columns O and P with the info you get from your scientist (these values will be used to calculate the soft limits in sardana from the hard limits position info you get during testing).
+ Modify 'Tests' sheet
- Add E column (Sardana axis name) (takes info from Data)
- Add columns S(I*R) to V(IVOLT configured).
- Be careful here in case you have put info in old AA (Homing position) by hand during your tests. Homing columns Z and AA have to be changed by new AD (Home origin)-AH(Home done). 'Fat home position' is the info for the fat. AF contains the ABSOFFSET param you configure. AG Home position has to contain the HOMEPOS parameter you configure in the home tab of icepapcms.
This aims to separate ABSOFFSET from the fat home position.
- Inserts columns BC (stroke from encoder counts)-BO(Soft lim-). This is information for sardana configuration. Sign has to be +1 or -1. Offset defaults to 0
- Insert column BQ (gets tc device from Data)
- Insert column CB with the text chain for the tango attributes
160113 | Julio Lidón-Simón | Motors sheet:
Added NVOLT and IVOLT again.
Added following motors
Oriental PK245DB
Oriental PK245MD15B
Oriental PK245M-02B unipolar 1 coil
Oriental PK266D14B
Oriental PK268D14B
Oriental PKP264MU20B-L unipolar 1 coil
161031 | Julio Lidón-Simon | Data sheet:
Step division column header renamed to Full step division factor
Rated current column renamed to sinusoidal peak current
Idle current renamed to recommended idle current
Motors:
Removed lock
Tests:
Added INOM and IIDLE columns
A number of changes in the Tests sheet… too long to describe here.
161111 | Gabriela Todorescu | Data sheet:
renaming TGT encoder to 1st encoder
1st encoder source and 2nd encoder source added
Tests sheet:
2nd encoder cable column added
columns CI to CW related to 2nd encoder added
161121 | Julio Lidón Simón | Motors sheet:
Added all 3 missing motors to the motor sheet of the excel
Oriental PK268M-01 half coil
Nanotech AS4118L1804-E
Nanotech AS5918L4204-E
Added Renishaw 32bits MSBMSKNB 1 (31b)
Changed unipolar 1 coil motors to half coil
Tests sheet:
Added a column for NRES in TESTS
Data sheet:
Added two columns for repeatability and resolution in Data
161206 | Julio Lidón
- Data.Repeatabilty header changed (AN)
- Data. Disable signals (AL) changed to: Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd
- Configurations. C1 changed to: Axis name (ccdb)
- Configurations. I column formula was wrong (taken from 2 rows below than expected (from row 4 for row 2).
- FATPOSITIONS column F formula was wrong (taken from 2 rows below than expected (from row 4 for row 2).
- Tests. Column D header changed to: Axis name (ccdb)
- Tests. Column AA header changed to: NRES (?MEAS R)
- Tests. Column AF header changed to: Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd
- Tests.AY. Inserted new column with header: Encoder noise pkpk (ects)
- Tests.BR. Formula changed to have into account UseEncoderSource to: =IF(BP17="FALSE";($BM17*($AQ17/$BL17)+$BQ17);($BM17*($AR17/$AN17)+$BQ17))
- Tests.BP. Formula changed to have into account UseEncoderSource to: =IF(BP17="FALSE";($BM17*($AS17/$BL17)+$BQ17);($BM17*($AT17/$AN17)+$BQ17))
- Tests.BP column header changed to: UseEncoderSourceFormula: TRUE/FALSE
- Tests CA column: Inserted new column with header: Warning temperature
- Tests CB column: Inserted new column with header: Alarm temperature
- Tests CG column: Inserted new column with header: Encoder noise pkpk (units). Formula: =$AY2/$AN2
- Tests CH column: Inserted new column with header:Deadband in (units). Formula:=$CF2/($AP2*$AM2)
- Tests CI column: Inserted new column with header: Aimed repeatability (units). Formula:=Data!$AN2
- Tests CJ column: Inserted new column with header: Aimed resolution (units). Formula:=Data!$AO2
- Motors. Phytron ZSS.52.200.2.5 parallel appeared twice
- Motors. Oriental PK243-01-SGXX series renamed to Oriental PK243-01 series
- Motors. Oriental PK243-02-SGXX series renamed to Oriental PK243-02 series
- Motors. Oriental PK245DA is renamed to Oriental PK245D (since DA and DB are the same motor with just single or double shaft).
- Motors. Oriental PK245MD15B renamed to Oriental PK245MD15 (since A or B stand for the double shaft version)
- Motors. Oriental Oriental PK245M-02B series renamed to Oriental PK245M-02 series (final B stands for double shaft version)
- Motors. Oriental PK245M-02B half coil renamed to Oriental PK245M-02 half coil
- Motors. Oriental PK245M-02 half coil rated current increased by a factor 1.4
- Motors. Oriental PK266D14B renamed to Oriental PK266D14
- Motors. Oriental PK264JD renamed to Oriental PK264JD
- Motors. Oriental PK267JA series renamed to Oriental PK267J series
- Motors. Added Oriental PK267JD series
- Motors. Oriental PK268D14B renamed to Oriental PK268D14
- Motors. Oriental PK268M-01 half coil rated current increased by a factor 1.4
- Motors. Oriental PK296-AE-SGxx series renamed to Oriental PK296AE-SGxx series
- Motors. Oriental PK296-AE-SGxx parallel renamed to Oriental PK296AE-SGxx parallel
- Data. Added column AP for settling time constraints for fast step scans
- Tests. Added column CP for settling time constraints for fast step scans
- Tests. Added column CQ for validation text with who and when
- Manufacturers info sheet updated with document: https://alfresco.maxlab.lu.se/share/page/site/maxiv/document-details?nodeRef=workspace://SpacesStore/7adbde15-a2e8-41ef-a236-9060bda8fe38#
- Motors. NVOLT has been limited to 48V in the excel (to avoid driving motors with low resistance, low inductance and high rated currents with short test cables on too much voltage, you can increase the voltage if you need it for speed reasons yourself)
170116 | Julio Lidón | - Tests. Column CS with tango attribs chain. CONCATENATE function does not keep the number formatting and creates problems with . and , as separator for decimals. The symbol & solves that problem.
Therefore the column now has the following formula:
="Step_per_unit:"&$BL2&";"&"Offset:"&$BQ2&";"&"Sign:"&$BM2&";"&"EncoderSource:"&$BN2&";"&"EncoderSourceFormula:"&$BO2&";"&"UseEncoderSource:"&$BP2
170130 | Gabriela Todorescu
Tango sheet
"device" -> "devicename"
"controller sardana" -> "controller"
"axis name sardana" -> "name"
"axis nr" -> "axis"
"attribute chain" -> "attributes"
"units" -> "unit"
Ready. (N or blank: don't take the data here)
Data sheet
column AB added; Motor mechanical resolution (full steps /2nd unit)
Tests sheet
column CY added; Motor resolution (fst per mm/deg)
column DC formula changed to =IF(K2="-";"-";(CZ2/CY2)*AO2); Theoretical 2nd encoder counts per motor turn

## Manufacturers info
Equiment | Component | Actuation | Motor model | Can the motor be moved independently from others (Yes/No) | Can the motor movement cause a collission (Yes/No) | Mechanical transmission details | Air/Vacuum | Step operation | Wiring (parallel/series) | Bipolar sinusoidal peak phase current (for specified wiring) | Rated idle Current | Linear motor mechanical resolution (units per full step) | Angular motor mechanical resolution if applies (specify unit) | Linear range (specify unit) | Angular range (specify unit) | Max speed (specify units) | Encoder Res. (unit/count) | Encoder Res (unit/count) if aplies | Encoder readhead, interpolator and scale model | Limit, home switch models | Home switch | Limits | Overtravel Limits | External disable connector | Motor direction (moves towards positive switch) | Home/nominal position (ects) | Limit- position (ects) | Limit+ position (ects) | Backlash | Parasitic effects | Thermocouples for EPS / Motor | Serial Device | Disable conditions | Settling time required (for fast step scans)
[31 empty] | - | - | - | 0.1s
[31 empty] | - | - | - | as fast as possible
[31 empty] | - | - | - | -
[31 empty] | 0/1 | - | Possible collision with axis X
[31 empty] | 0/1 | - | Thermocouple readout
[31 empty] | - | PI x 1 | -
[31 empty] | 0/1 | - | -
[31 empty] | 0/1 | - | -
[32 empty] | - | -
[32 empty] | - | -
[32 empty] | - | -
[32 empty] | - | -
[32 empty] | - | -
[32 empty] | - | -

## SystemInfo
Number of rows/cables: | 200
IcePAP system master name | B108A-D100830-CAB04-CTL-IPAP-01.maxiv.lu.se
IcePAP system master mac | 00:0C:C6:77:B2:98
IcePAP system master IP | 172.16.190.55
Number of chassis
Number of drivers
Alfresco link to cable drawings
Alfresco link to hardware overview
Alfresco link to motorization information excel
Labels printed
Client computer
Machine hosting icepapcms database
Sardana pool generated

## Data
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Axis name (sardana) | Motor model and wiring (=template) | Independent axis (Y/N) | Can axis collide (Y/N) | Mechanical transmission details | Unit | Air/Vacuum | Full step division factor | Sinusoidal peak current (A) | Resistance per phase (Ohm) | Inductance per phase(mH) | Motor full steps | Recommended idle current % of nominial | Motor mechanical resolution (full steps /unit) | Range (units) | Max speed(fst/s) | 1st encoder source (ABSENC, ENCIN, INPOS, NONE) | Theoretical 1st encoder counts per motor turn | 1st encoder Res. (counts/unit) | 1st encoder readhead, interpolator and scale model | 1st encoder template (if absolute) | Home origin (absolute, limit-, encoder, homesw) | 1st encoder supply (Voltage supply, signal levels (TTL/RS422) and current consumption) | Motor mechanical resolution (full steps /2nd unit) | 2nd encoder source (ABSENC, ENCIN, INPOS, NONE) | Theoretical 2nd encoder counts per motor turn | 2nd encoder Res. (counts/unit) | 2nd encoder readhead, interpolator and scale model | 2nd encoder template (if absolute) | 2nd encoder supply (Voltage supply, signal levels (TTL/RS422) and current consumption) | Limits ('-' if none) | Overtravel Limits ('-' if none) | Thermocouples for MPS / Motor ('-' if none) | Associated TC device | Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd | Limit switches type | Repeatability (units) | Resolution (units) | Settling time (if fast step scans needed) | Motor direction | Backlash | Parasitic effects | Serial Device | Camera | Mech. Drwg | Elec. Drwg
1 | B108A-FE-OPT-MM-01 | Movable mask unit | X | mm1_x | Oriental PK268D14 | Y | N | 1mm thread pitch | mm | Air | 2 | 2 | 4.5 | 14.4 | 200 | - | 200 | ±6mm |  | ABSENC | 1000000 | 1000000 | Renishaw Resolute RL32BAT001B30A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 | - | - | GND |  | rep | res | 0.1s
2 | B108A-FE-OPT-MM-01 | Movable mask unit | Y | mm1_y | Oriental PK268D14 | Y | N | 1:4 elevating stage,1mm thread pitch | mm | Air | 2 | 2 | 4.5 | 14.4 | 200 | - | 800 | ±6mm |  | ABSENC | 250000 | 1000000 | Renishaw Resolute RL32BAT001B30A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 | - | - | GND |  |  |  | lowest possible
3 | B108A-FE-OPT-MM-02 | Movable mask unit | X | mm2_x | Oriental PK268D14 | Y | N | 1mm thread pitch | mm | Air | 2 | 2 | 4.5 | 14.4 | 200 | - | 200 | ±6mm |  | ABSENC | 1000000 | 1000000 | Renishaw Resolute RL32BAT001B30A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 | - | - | GND
4 | B108A-FE-OPT-MM-02 | Movable mask unit | Y | mm2_y | Oriental PK268D14 | Y | N | 1:4 elevating stage,1mm thread pitch | mm | Air | 2 | 2 | 4.5 | 14.4 | 200 | - | 800 | ±6mm |  | ABSENC | 250000 | 1000000 | Renishaw Resolute RL32BAT001B30A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 0 | - | - | GND
5 | B108A-O01-OPT-MIR-01 | Mirror Species | VC | m1_vc | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
6 | B108A-O01-OPT-MIR-01 | Mirror Species | VR | m1_vr | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
7 | B108A-O01-OPT-MIR-01 | Mirror Species | VL | m1_vl | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
8 | B108A-O01-OPT-MIR-01 | Mirror Species | HLAT | m1_x | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
11 | B108A-O01-OPT-MIR-01 | Mirror Species | HPIT | m1_hpit | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle; Lever arm length 300mm | mrad | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 5882.3529 | ±3mm; ±10 mrad |  | ABSENC | 20029.4551 | 294550.81 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
12 | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PIT | mono_m2pit | Oriental PK268-03 series | Y | Y | Thread Pitch 2mm; Gear ratio 50:1 | mrad | Air | 16 | 2.9 | 2 | 6.4 | 200 | - | 1929.5 | 345.4147 |  | ENCIN | 950.228 | 9167.3247 | Heidenhain RON 905 UHV + Interface box to computer card | - | encoder | [8 empty] | 2 | 2 | - | - | E(2O+D)
13 | B108A-O02-OPT-MONO-01 | Monochromator Species | GPIT | mono_gpit | Oriental PK268-03 series | Y | Y | Thread Pitch 2mm; Gear ratio 50:1 | mrad | Air | 16 | 2.9 | 2 | 6.4 | 200 | - | 1929.5 | 420.6787 |  | ENCIN | 950.228 | 9167.3247 | Heidenhain RON 905 UHV + Interface box to computer card | - | encoder | [8 empty] | 2 | 2 | - | - | E(2O+D)
14 | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXC | mono_gexc | Oriental PK268M-E2.0 parallel | Y | Y | Thread Pitch 3mm; Gear ratio 1:1 | mm | Air | 2 | 3.9 | 1.13 | 4.8 | 400 | - | 133.3333 | 207.32 |  | ENCIN | 30000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | E(2O+D)
15 | B108A-O02-OPT-BAFF-01 | Baffle Species | VT | mono_baff_vt | Oriental PK268M-E2.0 parallel | Y | N | Gear ratio 2:1; 4mm Pitch | mm | Air | 2 | 3.9 | 1.13 | 4.8 | 400 | - | 200 | 46mm |  | ENCIN | 20000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Cherry D 4 3 V3 RA
16 | B108A-O02-OPT-BAFF-01 | Baffle Species | VB | mono_baff_vb | Oriental PK268M-E2.0 parallel | Y | N | Gear ratio 2:1; 4mm Pitch | mm | Air | 2 | 3.9 | 1.13 | 4.8 | 400 | - | 200 | 46mm |  | ENCIN | 20000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Cherry D 4 3 V3 RA
17 | B108A-O02-OPT-BAFF-01 | Baffle Species | HR | mono_baff_hr | Oriental PK268M-E2.0 parallel | Y | N | Gear ratio 2:1; 4mm Pitch | mm | Air | 2 | 3.9 | 1.13 | 4.8 | 400 | - | 200 | 46mm |  | ENCIN | 20000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Cherry D 4 3 V3 RA
18 | B108A-O02-OPT-BAFF-01 | Baffle Species | HL | mono_baff_hl | Oriental PK268M-E2.0 parallel | Y | N | Gear ratio 2:1; 4mm Pitch | mm | Air | 2 | 3.9 | 1.13 | 4.8 | 400 | - | 200 | 46mm |  | ENCIN | 20000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Cherry D 4 3 V3 RA
21 | B108A-O03-OPT-BAFF-01 | Baffle Species | VT | m3_baff_vt | Nanotec L4118S1404-M6X1 | Y | N | M6x 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | 20mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND
22 | B108A-O03-OPT-BAFF-01 | Baffle Species | VB | m3_baff_vb | Nanotec L4118S1404-M6X1 | Y | N | M6x 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | 20mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND
23 | B108A-O03-OPT-BAFF-01 | Baffle Species | HL | m3_baff_hl | Nanotec L4118S1404-M6X1 | Y | N | M6x 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | 20mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND
24 | B108A-O03-OPT-BAFF-01 | Baffle Species | HR | m3_baff_hr | Nanotec L4118S1404-M6X1 | Y | N | M6x 1mm thread pitch | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | 20mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND
25 | B108A-O03-OPT-MIR-03 | Mirror Species | VC | m3_vc | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
26 | B108A-O03-OPT-MIR-03 | Mirror Species | VR | m3_vr | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
27 | B108A-O03-OPT-MIR-03 | Mirror Species | VL | m3_vl | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
28 | B108A-O03-OPT-MIR-03 | Mirror Species | HLAT | m3_x | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
31 | B108A-O03-OPT-MIR-03 | Mirror Species | HPIT | m3_hpit | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle; Lever arm length 300mm | mrad | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 5882.3529 | ±3mm; ±10 mrad |  | ABSENC | 20029.4551 | 294550.81 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
32 | B108A-O03-OPT-MIR-03 | Mirror Species | MCHG | m3_mchg | Nanotec L4118S1404-M6X1 | Y | N |  | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±32mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
33 | B108A-OA04-DIA-BPM-01 | BPM Species | V | a_bpm | Wantai 42BYGHM809 | Y | N | Gear? | mm | Air | 2 | 2.4 | 1.8 | 2.8 | 400 | - | 24000 | 120mm |  | ABSENC | 333.3333 | 20000 | Renishaw RESOLUTE RL26BAE050D30F | Renishaw Resolute 26 bits BISS-C | absolute | [8 empty] | 2 | 0 | - | - | GND
34 | B108A-OB04-DIA-BPM-01 | BPM Species | V | b_bpm | Wantai 42BYGHM809 | Y | N | Gear? | mm | Air | 2 | 2.4 | 1.8 | 2.8 | 400 | - | 24000 | 120mm |  | ABSENC | 333.3333 | 20000 | Renishaw RESOLUTE RL26BAE050D30F | Renishaw Resolute 26 bits BISS-C | absolute | [8 empty] | 2 | 0 | - | - | GND
35 | B108A-OA04-OPT-SLIT-01 | Slits Species | V | a_slit1_v | Oriental PK244-01 parallel | Y | N | Harmonic Drive 50:1 Gearbox; Spindle M8x1mm pitch; In-vacuum Lever 2:1 | um | Air | 2 | 2.4 | 1.8 | 3.2 | 200 | - | 20 | 4996um; 2498um slit gap |  | ENCIN | 200 | 20 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Metrol precision switch PT5S1CB
36 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAR | a_slit1_hdar | McLennan 23HS-108 Mk2 | Y | N | Gear ratio 1:1; Spindle pitch 1mm | mm | Air | 16 | 5.5 | 0.37 | 0.59 | 200 | - | 200 | ±12.5mm |  | NONE | 0 |  | - | - | [9 empty] | 2 | 0 | - | - | GND
37 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAL | a_slit1_hdal | McLennan 23HS-108 Mk2 | Y | N | Gear ratio 1:1; Spindle pitch 1mm | mm | Air | 16 | 5.5 | 0.37 | 0.59 | 200 | - | 200 | ±12.5mm |  | NONE | 0 |  | - | - | [9 empty] | 2 | 0 | - | - | GND
38 | B108A-OA04-OPT-SLIT-01 | Slits Species | HZ | a_slit1_hz | Oriental PK268M-E2.0 parallel | Y | N | Gear ratio 1:1; Spindle pitch 2mm | mm | Air | 2 | 3.9 | 1.13 | 4.8 | 400 | - | 200 | ±50mm |  | ENCIN | 20000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Cherry D 4 3 V3 RA
41 | B108A-OB04-OPT-SLIT-01 | Slits Species | V | b_slit1_v | Oriental PK244-01 parallel | Y | N | Harmonic Drive 50:1 Gearbox; Spindle M8x1mm pitch; In-vacuum Lever 2:1 | um | Air | 2 | 2.4 | 1.8 | 3.2 | 200 | - | 20 | 4903um; 2451um slit gap |  | ENCIN | 200 | 20 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Metrol precision switch PT5S1CB
42 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAR | b_slit1_hdar | McLennan 23HS-108 Mk2 | Y | N | Gear ratio 1:1; Spindle pitch 1mm | mm | Air | 16 | 5.5 | 0.37 | 0.59 | 200 | - | 200 | ±12.5mm |  | NONE | 0 |  | - | - | [9 empty] | 2 | 0 | - | - | GND
43 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAL | b_slit1_hdal | McLennan 23HS-108 Mk2 | Y | N | Gear ratio 1:1; Spindle pitch 1mm | mm | Air | 16 | 5.5 | 0.37 | 0.59 | 200 | - | 200 | ±12.5mm |  | NONE | 0 |  | - | - | [9 empty] | 2 | 0 | - | - | GND
44 | B108A-OB04-OPT-SLIT-01 | Slits Species | HZ | b_slit1_hz | Oriental PK268M-E2.0 parallel | Y | N | Gear ratio 1:1; Spindle pitch 2mm | mm | Air | 2 | 3.9 | 1.13 | 4.8 | 400 | - | 200 | ±50mm |  | ENCIN | 20000 | 10000 | RenishawTONiC T1001-10A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 2 | - | - | 2O | Cherry D 4 3 V3 RA
45 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VT | a_m4_baff_vt | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
46 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VB | a_m4_baff_vb | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
47 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HL | a_m4_baff_hl | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
48 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HR | a_m4_baff_hr | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
51 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VT | b_m4_baff_vt | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
52 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VB | b_m4_baff_vb | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
53 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HL | b_m4_baff_hl | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
54 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HR | b_m4_baff_hr | Nanotec L4118S1404-M6X1 | Y | N | Spindle thread type and pitch M6x1mm | mm | Air | 2 | 2 | 2 | 2.8 | 200 | - | 200 | ±10mm |  | ENCIN | 10000 | 10000 | RenishawTONiC T1001-05A with Ti0200E04A interface | - | encoder | [8 empty] | 2 | 0 | - | - | GND |  | 10um | 5um
55 | B108A-OA05-OPT-MIR-04 | Mirror Species | VC | a_m4_vc | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
56 | B108A-OA05-OPT-MIR-04 | Mirror Species | VR | a_m4_vr | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
57 | B108A-OA05-OPT-MIR-04 | Mirror Species | VL | a_m4_vl | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
58 | B108A-OA05-OPT-MIR-04 | Mirror Species | HLAT | a_m4_x | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
61 | B108A-OA05-OPT-MIR-04 | Mirror Species | HPIT | a_m4_hpit | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mrad | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 5882.3529 | ±3mm; ±10 mrad |  | ABSENC | 20029.4551 | 294550.81 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
62 | B108A-OB05-OPT-MIR-04 | Mirror Species | VC | b_m4_vc | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
63 | B108A-OB05-OPT-MIR-04 | Mirror Species | VR | b_m4_vr | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
64 | B108A-OB05-OPT-MIR-04 | Mirror Species | VL | b_m4_vl | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
65 | B108A-OB05-OPT-MIR-04 | Mirror Species | HLAT | b_m4_x | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mm | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 20000 | ±10mm |  | ABSENC | 20000 | 1000000 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
66 | B108A-OB05-OPT-MIR-04 | Mirror Species | HPIT | b_m4_hpit | Oriental PK264M-E2.0 parallel | Y | N | Harmonic Drive HFUC-14-50-2UH-SP 50:1 Gearbox; M14x1mm spindle | mrad | Air | 2 | 3.9 | 0.7 | 1.7 | 400 | - | 5882.3529 | ±3mm; ±10 mrad |  | ABSENC | 20029.4551 | 294550.81 | Renishaw Resolute RL32BAT001B05A, RTLA-S scale | Renishaw Resolute 32 bits BISS-C | absolute | [8 empty] | 2 | 2 | - | - | 2O | Metrol PT5M1CB
67 | B108A-FE-DIA-SCRN-01 | Fluorescence Screen Front End 1.5GeV | Y |  |  | Y | N |  | mm | Air | 2 | #N/A | #N/A | #N/A | #N/A | - | [19 empty] | - | -
68 | [5 empty] | Y | N |  |  |  |  | #N/A | #N/A | #N/A | #N/A | - | [19 empty] | - | -

## Configurations
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Motor cable | Encoder cable ('-' if none) | ANSTEP | ANTURN | TGTENC | ABSNSTEP | ABSNTURN | EINNSTEP | EINNTURN
1 | B108A-FE-OPT-MM-01 | Movable mask unit | X | #N/A | #N/A | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
2 | B108A-FE-OPT-MM-01 | Movable mask unit | Y | #N/A | #N/A | 400 | 1 | ABSENC | 250000 | 1 |  | 1
3 | B108A-FE-OPT-MM-02 | Movable mask unit | X | #N/A | #N/A | 400 | 1 | ABSENC | 1000000 | 1 |  | 1
4 | B108A-FE-OPT-MM-02 | Movable mask unit | Y | #N/A | #N/A | 400 | 1 | ABSENC | 250000 | 1 |  | 1
5 | B108A-O01-OPT-MIR-01 | Mirror Species | VC | W041395 | W041396 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
6 | B108A-O01-OPT-MIR-01 | Mirror Species | VR | W041397 | W041398 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
7 | B108A-O01-OPT-MIR-01 | Mirror Species | VL | W041399 | W041400 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
8 | B108A-O01-OPT-MIR-01 | Mirror Species | HLAT | W041401 | W041402 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
0
11 | B108A-O01-OPT-MIR-01 | Mirror Species | HPIT | W041403 | W041404 | 800 | 1 | ABSENC | 20029.4551 | 1 |  | 1
12 | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PIT | W041405 | W041406 | 3200 | 1 | ENCIN |  | 1 | 950.228 | 1
13 | B108A-O02-OPT-MONO-01 | Monochromator Species | GPIT | W041407 | W041408 | 3200 | 1 | ENCIN |  | 1 | 950.228 | 1
14 | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXC | W041409 | W041410 | 800 | 1 | ENCIN |  | 1 | 30000 | 1
15 | B108A-O02-OPT-BAFF-01 | Baffle Species | VT | W041411 | W041412 | 800 | 1 | ENCIN |  | 1 | 20000 | 1
16 | B108A-O02-OPT-BAFF-01 | Baffle Species | VB | W041413 | W041414 | 800 | 1 | ENCIN |  | 1 | 20000 | 1
17 | B108A-O02-OPT-BAFF-01 | Baffle Species | HR | W041415 | W041416 | 800 | 1 | ENCIN |  | 1 | 20000 | 1
18 | B108A-O02-OPT-BAFF-01 | Baffle Species | HL | W041417 | W041418 | 800 | 1 | ENCIN |  | 1 | 20000 | 1
0
0
21 | B108A-O03-OPT-BAFF-01 | Baffle Species | VT | W041419 | W041420 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
22 | B108A-O03-OPT-BAFF-01 | Baffle Species | VB | W041421 | W041422 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
23 | B108A-O03-OPT-BAFF-01 | Baffle Species | HL | W041423 | W041424 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
24 | B108A-O03-OPT-BAFF-01 | Baffle Species | HR | W041425 | W041426 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
25 | B108A-O03-OPT-MIR-03 | Mirror Species | VC | W041427 | W041428 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
26 | B108A-O03-OPT-MIR-03 | Mirror Species | VR | W041429 | W041430 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
27 | B108A-O03-OPT-MIR-03 | Mirror Species | VL | W041431 | W041432 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
28 | B108A-O03-OPT-MIR-03 | Mirror Species | HLAT | W041433 | W041434 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
0
0
31 | B108A-O03-OPT-MIR-03 | Mirror Species | HPIT | W041435 | W041436 | 800 | 1 | ABSENC | 20029.4551 | 1 |  | 1
32 | B108A-O03-OPT-MIR-03 | Mirror Species | MCHG | W041437 | W041438 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
33 | B108A-OA04-DIA-BPM-01 | BPM Species | V | W041439 | W041440 | 800 | 1 | ABSENC | 333.3333 | 1 |  | 1
34 | B108A-OB04-DIA-BPM-01 | BPM Species | V | W041441 | W041442 | 800 | 1 | ABSENC | 333.3333 | 1 |  | 1
35 | B108A-OA04-OPT-SLIT-01 | Slits Species | V | W041443 | W041444 | 400 | 1 | ENCIN |  | 1 | 200 | 1
36 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAR | W041445 | #N/A | 3200 | 1 | NONE |  | 1 |  | 1
37 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAL | W041446 | #N/A | 3200 | 1 | NONE |  | 1 |  | 1
38 | B108A-OA04-OPT-SLIT-01 | Slits Species | HZ | W041447 | W041448 | 800 | 1 | ENCIN |  | 1 | 20000 | 1
0
0
41 | B108A-OB04-OPT-SLIT-01 | Slits Species | V | W041449 | W041450 | 400 | 1 | ENCIN |  | 1 | 200 | 1
42 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAR | W041451 | #N/A | 3200 | 1 | NONE |  | 1 |  | 1
43 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAL | W041452 | #N/A | 3200 | 1 | NONE |  | 1 |  | 1
44 | B108A-OB04-OPT-SLIT-01 | Slits Species | HZ | W041453 | W041454 | 800 | 1 | ENCIN |  | 1 | 20000 | 1
45 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VT | W041455 | W041456 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
46 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VB | W041457 | W041458 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
47 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HL | W041459 | W041460 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
48 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HR | W041461 | W041462 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
0
0
51 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VT | W041463 | W041464 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
52 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VB | W041465 | W041466 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
53 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HL | W041467 | W041468 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
54 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HR | W041469 | W041470 | 400 | 1 | ENCIN |  | 1 | 10000 | 1
55 | B108A-OA05-OPT-MIR-04 | Mirror Species | VC | W041471 | W041472 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
56 | B108A-OA05-OPT-MIR-04 | Mirror Species | VR | W041473 | W041474 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
57 | B108A-OA05-OPT-MIR-04 | Mirror Species | VL | W041475 | W041476 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
58 | B108A-OA05-OPT-MIR-04 | Mirror Species | HLAT | W041477 | W041478 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
0
0
61 | B108A-OA05-OPT-MIR-04 | Mirror Species | HPIT | W041479 | W041480 | 800 | 1 | ABSENC | 20029.4551 | 1 |  | 1
62 | B108A-OB05-OPT-MIR-04 | Mirror Species | VC | W041481 | W041482 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
63 | B108A-OB05-OPT-MIR-04 | Mirror Species | VR | W041483 | W041484 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
64 | B108A-OB05-OPT-MIR-04 | Mirror Species | VL | W041485 | W041486 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
65 | B108A-OB05-OPT-MIR-04 | Mirror Species | HLAT | W041487 | W041488 | 800 | 1 | ABSENC | 20000 | 1 |  | 1
66 | B108A-OB05-OPT-MIR-04 | Mirror Species | HPIT | W041489 | W041490 | 800 | 1 | ABSENC | 20029.4551 | 1 |  | 1
67 | B108A-FE-DIA-SCRN-01 | Fluorescence Screen Front End 1.5GeV | Y | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
68 | 0 | 0 | 0 | #N/A | #N/A | #N/A | 1 | 0 |  | 1 |  | 1
0
0

## FATPOSITIONS
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Units | Encoder input | Nominal position (ects) | FAT Home mark (ects) | FAT LIM- (ects) | FAT LIM+ (ects) | FAT Home (units) | FAT LIM- (units) | FAT LIM+ (units) | Margin soft limit- (units) | Margin soft limit+ (units)
1 | B108A-FE-OPT-MM-01 | Movable mask unit | X | mm | ABSENC
2 | B108A-FE-OPT-MM-01 | Movable mask unit | Y | mm | ABSENC
3 | B108A-FE-OPT-MM-02 | Movable mask unit | X | mm | ABSENC
4 | B108A-FE-OPT-MM-02 | Movable mask unit | Y | mm | ABSENC
5 | B108A-O01-OPT-MIR-01 | Mirror Species | VC | mm | ABSENC |  | 30555340 | 19885810 | 41188700 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M1-drives; Vertical linear drive; ball-cone
6 | B108A-O01-OPT-MIR-01 | Mirror Species | VR | mm | ABSENC |  | 31047130 | 20368710 | 41685230 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M1-drives; Vertical linear drive; flat surface
7 | B108A-O01-OPT-MIR-01 | Mirror Species | VL | mm | ABSENC |  | 30673040 | 20029280 | 41319500 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M1-drives; Vertical linear drive; V-groove
8 | B108A-O01-OPT-MIR-01 | Mirror Species | HLAT | mm | ABSENC |  | 45461460 | 34795320 | 56083450 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M1-drives; Horizontal lateral translation
0 | 0 | 0 | 0 | 0 | 0
11 | B108A-O01-OPT-MIR-01 | Mirror Species | HPIT | mrad | ABSENC |  | 35277620 | 3215148 | 38432200 | 0 | ~-3.5 | ~3.5 |  |  | 20120117-FAT-M1-drives; Pitch angular drive
12 | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PIT | mrad | ENCIN | [5 empty] | -325471.8351 | 19810.1367 |  |  | 20120117_Species_PGM_Mirror_Pitch_01
13 | B108A-O02-OPT-MONO-01 | Monochromator Species | GPIT | mrad | ENCIN | [5 empty] | -40590.3128 | 15346.8643 |  |  | 20120117_Species_PGM_grating_Pitch_01
14 | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXC | mm | ENCIN | [5 empty] | -181.311 | 26.016 |  |  | 20120117_Species_PGM_Linear_Grating_01
15 | B108A-O02-OPT-BAFF-01 | Baffle Species | VT | mm | ENCIN |  | 7365 | -223630 | 229130 | 0.7365 | -22.363 | 22.913 |  |  | 20120117_V-OLCL-RB-Linear-MVRMS_top
16 | B108A-O02-OPT-BAFF-01 | Baffle Species | VB | mm | ENCIN |  |  | -229830 | 229950 |  | -22.983 | 22.995 |  |  | 20120117_V-OLCL-RB-Linear-MVRMS_bottom
17 | B108A-O02-OPT-BAFF-01 | Baffle Species | HR | mm | ENCIN |  | 3242 | -227490 | 228360 | 0.3242 | -22.749 | 22.836 |  |  | 20120117_V-OLCL-RB-Linear-MVRMS_right
18 | B108A-O02-OPT-BAFF-01 | Baffle Species | HL | mm | ENCIN |  | 2744 | -230880 | 229170 | 0.2744 | -23.088 | 22.917 |  |  | 20120117_V-OLCL-RB-Linear-MVRMS_top
0 | 0 | 0 | 0 |  | 0
0 | 0 | 0 | 0 |  | 0
21 | B108A-O03-OPT-BAFF-01 | Baffle Species | VT | mm | ENCIN |  | -11707 | -112123 | 88140 | 0 | -10.04 | 9.98 |  |  | W001017-03 Slits Test Record; Vertical Top Blade
22 | B108A-O03-OPT-BAFF-01 | Baffle Species | VB | mm | ENCIN |  | -6628 | -106760 | 93339 | 0 | -10.01 | 10 |  |  | W001017-03 Slits Test Record; Vertical Bottom Blade
23 | B108A-O03-OPT-BAFF-01 | Baffle Species | HL | mm | ENCIN |  | -1381 | -101433 | 98457 | 0 | -10.01 | 9.98 |  |  | W001017-03 Slits Test Record; Horizontal Left Blade
24 | B108A-O03-OPT-BAFF-01 | Baffle Species | HR | mm | ENCIN |  | -13606 | -113511 | 86648 | 0 | -9.99 | 10.03 |  |  | W001017-03 Slits Test Record; Horizontal Right Blade
25 | B108A-O03-OPT-MIR-03 | Mirror Species | VC | mm | ABSENC |  | 30254060 | 19604250 | 40859680 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M3-drives; Vertical linear drive; ball-cone
26 | B108A-O03-OPT-MIR-03 | Mirror Species | VR | mm | ABSENC |  | 30657070 | 19834590 | 41178010 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M3-drives; Vertical linear drive; flat surface
27 | B108A-O03-OPT-MIR-03 | Mirror Species | VL | mm | ABSENC |  | 31187290 | 20541080 | 40806980 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M3-drives; Vertical linear drive; V-groove
28 | B108A-O03-OPT-MIR-03 | Mirror Species | HLAT | mm | ABSENC |  | 45348700 | 34695970 | 56029340 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M3-drives; Horizontal lateral translation
0 | 0 | 0 | 0 |  | 0
0 | 0 | 0 | 0 |  | 0
31 | B108A-O03-OPT-MIR-03 | Mirror Species | HPIT | mrad | ABSENC |  | 35612000 | 32274810 | 38734300 | 0 | ~-3.5 | ~3.5 |  |  | 20120117-FAT-M3-drives; Pitch angular drive
32 | B108A-O03-OPT-MIR-03 | Mirror Species | MCHG | mm | ENCIN |  | 105974 | -66549 | 278450 | 0 | ~-16 | ~16 |  |  | 20120117-FAT-M3-drives; Horizontal mirror switching drive
33 | B108A-OA04-DIA-BPM-01 | BPM Species | V | mm | ABSENC
34 | B108A-OB04-DIA-BPM-01 | BPM Species | V | mm | ABSENC
35 | B108A-OA04-OPT-SLIT-01 | Slits Species | V | um | ENCIN |  | 0 | -20448 | 29507 | 1365 | -110.4 | 2387.4 |  |  | 20120117_S008833_Xslit_348-08-06-00_17_11points_5cycles_CL_2013-05-13
36 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAR | mm | NONE |  | 0 | -18862.4 | 19096.8 | 13.9 | 19.8 | 7.9 |  |  | 20120117_S008833_Exit-slit-preblades-longitudinal; Right Blade
37 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAL | mm | NONE |  | 0 | -18467.2 | 19145.6 | 13.5 | 19.2 | 7.5 |  |  | 20120117_S008833_Exit-slit-preblades-longitudinal; Left Blade
38 | B108A-OA04-OPT-SLIT-01 | Slits Species | HZ | mm | ENCIN |  | -0.324 | -487370 | 487370 | 0 | ~-45 | ~45 |  |  | 20120117_S008833_Exit-slit-preblades-longitudinal; Longitudinal stage
41 | B108A-OB04-OPT-SLIT-01 | Slits Species | V | um | ENCIN |  | 0 | -19736 | 29292 | 1398.3 | -66.4 | 2385.1
42 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAR | mm | NONE |  | 0 | -19306.4 | 19316.8 | 13.9 | 19.9 | 7.9 |  |  | 20120117_S008834_Exit-slit-preblades-longitudinal; Right Blade
43 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAL | mm | NONE |  | 0 | -18134.4 | 18988 | 13.2 | 18.9 | 7.3 |  |  | 20120117_S008834_Exit-slit-preblades-longitudinal; Left Blade
44 | B108A-OB04-OPT-SLIT-01 | Slits Species | HZ | mm | ENCIN |  | -0.2422 | -48582 | 478266 | 0 | ~-45 | ~45 |  |  | 20120117_S008834_Exit-slit-preblades-longitudinal; Longitudinal stage
45 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VT | mm | ENCIN |  | 2065 | -98155 | 102132 | 0 | -10.02 | 10.01 |  |  | W001017-02 Slits Test Record; Vertical Top Blade
46 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VB | mm | ENCIN |  | 7250 | -92556 | 107334 | 0 | -9.98 | 10.01 |  |  | W001017-02 Slits Test Record; Vertical Bottom Blade
47 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HL | mm | ENCIN |  | -15490 | -115478 | 84576 | 0 | -10 | 10.01 |  |  | W001017-02 Slits Test Record; Horizontal Left Blade
48 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HR | mm | ENCIN |  | 13248 | -86695 | 112747 | 0 | -9.99 | 9.95 |  |  | W001017-02 Slits Test Record; Horizontal Right Blade
51 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VT | mm | ENCIN |  | 8984 | -90968 | 108995 | 0 | -10 | 10 |  |  | W001017-01 Slits Test Record; Vertical Top Blade
52 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VB | mm | ENCIN |  | -3793 | -103536 | 96474 | 0 | -9.97 | 10.03 |  |  | W001017-01 Slits Test Record; Vertical Bottom Blade
53 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HL | mm | ENCIN |  | 6826 | -93415 | 106568 | 0 | -10.02 | 9.97 |  |  | W001017-01 Slits Test Record; Horizontal Left Blade
54 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HR | mm | ENCIN |  | 9685 | -90340 | 109821 | 0 | -10 | 10.01 |  |  | W001017-01 Slits Test Record; Horizontal Right Blade
55 | B108A-OA05-OPT-MIR-04 | Mirror Species | VC | mm | ABSENC |  | 30511950 | 19864150 | 41147710 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-HP-drives; Vertical linear drive; ball-cone
56 | B108A-OA05-OPT-MIR-04 | Mirror Species | VR | mm | ABSENC |  | 30028150 | 19441640 | 40571580 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-HP-drives; Vertical linear drive; flat surface
57 | B108A-OA05-OPT-MIR-04 | Mirror Species | VL | mm | ABSENC |  | 30551700 | 19848610 | 41080910 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-HP-drives; Vertical linear drive; V-groove
58 | B108A-OA05-OPT-MIR-04 | Mirror Species | HLAT | mm | ABSENC |  | 45038630 | 34402140 | 55824330 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-HP-drives; Horizontal lateral translation
61 | B108A-OA05-OPT-MIR-04 | Mirror Species | HPIT | mrad | ABSENC |  | 35837070 | 32778150 | 38942390 | 0 | ~-3.5 | ~3.5 |  |  | 20120117-FAT-M4-HP-drives; Pitch angular drive
62 | B108A-OB05-OPT-MIR-04 | Mirror Species | VC | mm | ABSENC |  | 30239670 | 19580510 | 40900370 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-RIXS-drives; Vertical linear drive; ball-cone
63 | B108A-OB05-OPT-MIR-04 | Mirror Species | VR | mm | ABSENC |  | 30361730 | 19719320 | 40947080 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-RIXS-drives; Vertical linear drive; flat surface
64 | B108A-OB05-OPT-MIR-04 | Mirror Species | VL | mm | ABSENC |  | 30286210 | 19625100 | 40805940 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-RIXS-drives; Vertical linear drive; V-groove
65 | B108A-OB05-OPT-MIR-04 | Mirror Species | HLAT | mm | ABSENC |  | 45573530 | 34978650 | 56166560 | 0 | ~-10.5 | ~10.5 |  |  | 20120117-FAT-M4-RIXS-drives; Horizontal lateral translation
66 | B108A-OB05-OPT-MIR-04 | Mirror Species | HPIT | mrad | ABSENC |  | 35812410 | 32694310 | 38960370 | 0 | ~-3.5 | ~3.5 |  |  | 20120117-FAT-M4-RIXS-drives; Pitch angular drive
67 | B108A-FE-DIA-SCRN-01 | Fluorescence Screen Front End 1.5GeV | Y | mm | 0
68 | 0 | 0 | 0 | 0 | 0
Note on I36: motor steps
Note on J36: motor steps
Note on I37: motor steps
Note on J37: motor steps
Note on I42: motor steps
Note on J42: motor steps
Note on I43: motor steps
Note on J43: motor steps

## Tests
Axis nr | Equipment code | Equiment name | Axis name (ccdb) | Axis name (sardana) | Independent axis (Y/N) | Can axis collide (Y/N) | Disable (external 2xovertravel + disable ; open;gnd) | Motor cable | 1st encoder cable ('-' if none) | 2nd encoder cable ('-' if none) | Disable cable ('-' if none) | Cables labling checked | Motor template | Encoder template | Motor phase mH | Motor phase Ohms | measured mH | measured Ohms | R Cable (Ohms) | I*R | Vinductance (per 1khsts) | NVOLT | IVOLT | NCURR | ICURR | NRES (?MEAS R) | Lim- polarity (NORMAL/INVERTED) | Lim+ polarity(NORMAL/INVERTED) | Motor direction (NORMAL/INVERTED) | Encoder direction (NORMAL/INVERTED) | Disable signals: External_PLC 2xOvertravel + Disable + TC: E(2O+D+TC); Internal 2xOvertravel: 2O; open;gnd | Disable tests | Home origin (absolute, limit-, encoder, homesw) | FAT Home position | Absolute encoder offset (ects) (0 if no absolute encoder) | Home position (steps) | Home done | Motor resolution (fst per mm/deg) | Encoder Res. (counts/mm or deg) | Motor steps | ustep factor | lim- steps (releasing switch) | lim- ects ('-' if none) (releasing switch) | lim+ steps (releasing switch) | lim+ects (releasing switch) | Lim- position (ects or st) from FAT | Lim+ position (ects or st) from FAT | Lim- position (ects prehoming) to compare with FAT | Lim+ position (ects prehoming) to compare with FAT | Encoder noise pkpk (ects) | Speed | Acctime | Time between limits (s) | Theoretical encoder counts per motor turn | encoder counts per motor turn |  | Range (units) | stroke (from motor steps) in mm or deg | stroke (from encoder counts) in mm or deg | Units | Steps per unit (as calculated from motor steps) | Steps per unit (as calculated from encoder) | Tango Step_per_unit | Tango motor Sign(+1 or -1) | EncoderSource | EncoderSourceFormula | UseEncoderSourceFormula: TRUE/FALSE | Tango motor Offset (units) | Lim- position (units) | Lim+ position (units) | FAT lim- position (units) | FAT lim+ position (units) | Soft lim- (units) | Soft lim+ (units) | TC(MPS/Motor) | Associated TC device | In vacuum motor temperature after 5min operation (wherever it applies, otherwise '-') | Warning temperature | Alarm temperature | Disable cable and conditions tested (without plc, with final plc, final setup…) | Disable config | CL time constant | CL deadband | Encoder noise pkpk (units) | Deadband in (units) | Aimed repeatability (units) | Aimed resolution (units) | CL error allowed | Start velocity | Settle window (steps) | Settle time(s) | Fast settling time required for fast scan | Measured settling time (s) | Validation (who, when) | Comments | Tango attribs chain | 2nd encoder template | 2nd Encoder direction (NORMAL/INVERTED) | FAT Home position | Home position (steps) | Home done | Motor resolution (fst per mm/deg) | 2nd encoder Res. (counts/unit) | lim- 2nd ects ('-' if none) (releasing switch) | lim+ 2nd ects (releasing switch) | Theoretical 2nd encoder counts per motor turn | 2nd encoder counts per motor turn | 2nd encoder counts per motor turn configured | Range (units) | stroke (from 2nd encoder counts) (units) | Units | Steps per unit (as calculated from 2nd encoder)
1 | B108A-FE-OPT-MM-01 | Movable mask unit | X | mm1_x | Y | N | GND | W041387 | W041388 |  | - | ok | Oriental PK268D14 | Renishaw Resolute 32 bits BISS-C | 14.4 | 4.5 | 14.6 | 5.6 | 1.1 | 11.2 | 22.922 | 48 | 12 | 1.4 | 50 | 6.8 | N | N | N | N | GND | N/A | absolute | 0 | 0 | 48161702 | N/A | 200 | 1000000 | 200 | 2 | 17098 | 42738218 | 21486 | 53705017 | 0 | 0 | 42738218 | 53705017 |  | 400 | 0.15 | 10.97 | 1000000 | 999708.2042 | 1000000 | ±6mm | 10.97 | 10.9668 | mm | 400 | 400 | 400 | -1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | 48.1617 | 5.4235 | -5.5433 | 0 | 0 | 5.4235 | -5.5433 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | rep | res |  |  |  |  | 0.1s |  | Gabriela Todorescu; 2017-08-30 | FE-B-S3; | Step_per_unit:400;Offset:48.161702;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±6mm |  | mm | 0
2 | B108A-FE-OPT-MM-01 | Movable mask unit | Y | mm1_y | Y | N | GND | W041389 | W041390 |  | - | ok | Oriental PK268D14 | Renishaw Resolute 32 bits BISS-C | 14.4 | 4.5 | 14.8 | 5.5 | 1 | 11 | 23.236 | 48 | 12 | 1.4 | 50 | 7 | N | N | N | N | GND | N/A | absolute | 0 | 0 | 48410419 | N/A | 800 | 1000000 | 200 | 2 | 69136 | 43314212 | 85469 | 53527401 | 0 | 0 | 43314212 | 53527401 |  | 400 | 0.15 | 40.8325 | 250000 | 250124.0189 | 250000 | ±6mm | 10.2081 | 10.2132 | mm | 1600 | 1600 | 1600 | -1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | 48.4104 | 5.0962 | -5.117 | 0 | 0 | 5.0962 | -5.117 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | lowest possible |  | Gabriela Todorescu; 2017-08-30 | FE-B-S3; | Step_per_unit:1600;Offset:48.410419;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±6mm |  | mm | 0
3 | B108A-FE-OPT-MM-02 | Movable mask unit | X | mm2_x | Y | N | GND | W041391 | W041392 |  | - | ok | Oriental PK268D14 | Renishaw Resolute 32 bits BISS-C | 14.4 | 4.5 | 15 | 5.6 | 1.1 | 11.2 | 23.55 | 48 | 12 | 1.4 | 50 | 6.8 | N | N | N | N | GND | N/A | absolute | 0 | 0 | 48999867 | N/A | 200 | 1000000 | 200 | 2 | 17460 | 43651732 | 21817 | 54542930 | 0 | 0 | 43651732 | 54542930 |  | 400 | 0.15 | 10.8925 | 1000000 | 999880.4682 | 1000000 | ±6mm | 10.8925 | 10.8912 | mm | 400 | 400 | 400 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -48.9999 | -5.3481 | 5.5431 | 0 | 0 | -5.3481 | 5.5431 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-08-30 | FE-B-S3; | Step_per_unit:400;Offset:-48.999867;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±6mm |  | mm | 0
4 | B108A-FE-OPT-MM-02 | Movable mask unit | Y | mm2_y | Y | N | GND | W041393 | W041394 |  | - | ok | Oriental PK268D14 | Renishaw Resolute 32 bits BISS-C | 14.4 | 4.5 | 15 | 5.5 | 1 | 11 | 23.55 | 48 | 12 | 1.4 | 50 | 6.7 | N | N | N | N | GND | N/A | absolute | 0 | 0 | 48306431 | N/A | 800 | 1000000 | 200 | 2 | 69585 | 43382175 | 85585 | 53394599 | 0 | 0 | 43382175 | 53394599 |  | 400 | 0.15 | 40 | 250000 | 250310.6 | 250000 | ±6mm | 10 | 10.0124 | mm | 1600 | 1600 | 1600 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -48.3064 | -4.9243 | 5.0882 | 0 | 0 | -4.9243 | 5.0882 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-08-30 | FE-B-S3; | Step_per_unit:1600;Offset:-48.306431;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±6mm |  | mm | 0
5 | B108A-O01-OPT-MIR-01 | Mirror Species | VC | m1_vc | Y | N | 2O | W041395 | W041396 |  | - |  | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.5 | 1.8 | 1.1 | 7.02 | 4.5922 | 48 | 12 | 2 | 30 | 2.5 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30555340 | 0 | 32419926 | N/A | 20000 | 1000000 | 400 | 2 | 1096766 | 27418076 | 1496791 | 37420142 | 19885810 | 41188700 | 27418076 | 37420142 |  | 5000 | 0.3 | 80.005 | 20000 | 20002.8818 | 20000 | ±10mm | 10.0006 | 10.0021 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -32.4199 | -5.0018 | 5.0002 | ~-10.5 | ~10.5 | -5.0018 | 5.0002 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2019-03-13 | it's not possible to move from limit to limit because the bellows might get damaged | Step_per_unit:40000;Offset:-32.419926;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30555340 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
6 | B108A-O01-OPT-MIR-01 | Mirror Species | VR | m1_vr | Y | N | 2O | W041397 | W041398 |  | - |  | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.5 | 1.8 | 1.1 | 7.02 | 4.5922 | 48 | 12 | 2 | 30 | 2.5 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 31047130 | 0 | 31990461 | N/A | 20000 | 1000000 | 400 | 2 | 1079699 | 26992318 | 1479693 | 36993303 | 20368710 | 41685230 | 26992318 | 36993303 |  | 5000 | 0.3 | 79.9988 | 20000 | 20002.27 | 20000 | ±10mm | 9.9999 | 10.001 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -31.9905 | -4.9981 | 5.0028 | ~-10.5 | ~10.5 | -4.9981 | 5.0028 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2019-03-13 | it's not possible to move from limit to limit because the bellows might get damaged | Step_per_unit:40000;Offset:-31.990461;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 31047130 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
7 | B108A-O01-OPT-MIR-01 | Mirror Species | VL | m1_vl | Y | N | 2O | W041399 | W041400 |  | - |  | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.8 | 1.1 | 7.02 | 4.2861 | 48 | 12 | 2 | 30 | 2.5 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30673040 | 0 | 33327755 | N/A | 20000 | 1000000 | 400 | 2 | 1133091 | 28327381 | 1533119 | 38327947 | 20029280 | 41319500 | 28327381 | 38327947 |  | 5000 | 0.3 | 80.0056 | 20000 | 19999.732 | 20000 | ±10mm | 10.0007 | 10.0006 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -33.3278 | -5.0004 | 5.0002 | ~-10.5 | ~10.5 | -5.0004 | 5.0002 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2014-03-13 | it's not possible to move from limit to limit because the bellows might get damaged | Step_per_unit:40000;Offset:-33.327755;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30673040 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
8 | B108A-O01-OPT-MIR-01 | Mirror Species | HLAT | m1_x | Y | N | 2O | W041401 | W041402 |  | - |  | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.8 | 1.1 | 7.02 | 4.2861 | 48 | 12 | 2 | 30 | 2.5 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 45461460 | 0 | 45456340 | N/A | 20000 | 1000000 | 400 | 2 | 1618796 | 40464592 | 2018967 | 50462338 | 34795320 | 56083450 | 40464592 | 50462338 |  | 5000 | 0.3 | 80.0342 | 20000 | 19986.9476 | 20000 | ±10mm | 10.0043 | 9.9977 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -45.4615 | -4.9969 | 5.0009 | ~-10.5 | ~10.5 | -4.9969 | 5.0009 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2019-03-13 | it's not possible to move from limit to limit because the bellows might get damaged | Step_per_unit:40000;Offset:-45.46146;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 45461460 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
11 | B108A-O01-OPT-MIR-01 | Mirror Species | HPIT | m1_hpit | Y | N | 2O | W041403 | W041404 |  | - |  | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.8 | 1.1 | 7.02 | 4.2861 | 48 | 12 | 2 | 30 | 2.5 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 35277620 | 0 | 35212749 | N/A | 5882.3529 | 294550.81 | 400 | 2 | 1285644 | 32142689 | 1536922 | 38421261 | 3215148 | 38432200 | 32142689 | 38421261 |  | 2000 | 0.3 | 125.639 | 20029.4551 | 19989.2454 | 20000 | ±3mm; ±10 mrad | 21.3586 | 21.3158 | mrad | 11764.7059 | 11782.0324 | 11782 | 1 | attr://EncAbsEnc | VALUE/294550.81 | TRUE | -119.5473 | -10.4229 | 10.8929 | ~-3.5 | ~3.5 | -10.4229 | 10.8929 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2019-03-13 | it's not possible to move from limit to limit because the bellows might get damaged | Step_per_unit:11782;Offset:-119.547282860977;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | 0 |  | 35277620 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±3mm; ±10 mrad |  | mrad | 0
12 | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PIT | mono_m2pit | Y | Y | E(2O+D) | W041405 | W041406 |  | - |  | Oriental PK268-03 series | - | 6.4 | 2 | 0.85 | 2.8 | 0.8 | 8.12 | 1.935 | 48 | 12 | 2 | 50 | 3.5 | N | N | INV | N | E(2O+D) | OT+ ok; OT- ok | encoder | 0 | 0 | 0 | IK220 | 1929.5 | 9167.3247 | 200 | 16 | -10130197 | -3007400 | 518448 | 153800 | 0 | 0 | -3007400 | 153800 |  | 43000 | 2.68 | 247.6429 | 950.228 | 949.965 | 950 | 345.4147 | 344.9289 | 344.8334 | mrad | 30872 | 30879.4096 | 30879.44 | 1 | tango://b-v-species-csdb-0:10000/B108A/CTL/IK220-02/AnglesMean | ((VALUE*3.141559265)/180)*1000 | TRUE | 0 | -328.0564 | 16.777 | -325471.8351 | 19810.1367 | -328.0564 | 16.777 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2019-03-08 |  | Step_per_unit:30879.44;Offset:0;Sign:1;EncoderSource:tango://b-v-species-csdb-0:10000/B108A/CTL/IK220-02/AnglesMean;EncoderSourceFormula:((VALUE*3.141559265)/180)*1000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 345.4147 |  | mrad | 0
13 | B108A-O02-OPT-MONO-01 | Monochromator Species | GPIT | mono_gpit | Y | Y | E(2O+D) | W041407 | W041408 |  | - |  | Oriental PK268-03 series | - | 6.4 | 2 | 0.85 | 2.8 | 0.8 | 8.12 | 1.935 | 48 | 12 | 2 | 50 | 3.5 | N | N | INV | INV | E(2O+D) | OT+ ok; OT- ok | encoder | 0 | 0 | 0 | Ik220 | 1929.5 | 9167.3247 | 200 | 16 | -12468008 | -3701441 | 533299 | 158323 | 0 | 0 | -3701441 | 158323 |  | 60000 | 3.75 | 216.6884 | 950.228 | 950.0002 | 950 | 420.6787 | 421.1359 | 421.0349 | mrad | 30872 | 30879.4096 | 30879.44 | 1 | tango://b-v-species-csdb-0:10000/B108A/CTL/IK220-01/AnglesMean | ((VALUE*3.141559265)/180)*1000 | TRUE | 0 | -403.7646 | 17.2704 | -40590.3128 | 15346.8643 | -403.7646 | 17.2704 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2019-03-08 |  | Step_per_unit:30879.44;Offset:0;Sign:1;EncoderSource:tango://b-v-species-csdb-0:10000/B108A/CTL/IK220-01/AnglesMean;EncoderSourceFormula:((VALUE*3.141559265)/180)*1000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 420.6787 |  | mrad | 0
14 | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXC | mono_gexc | Y | Y | E(2O+D) | W041409 | W041410 |  | - |  | Oriental PK268M-E2.0 parallel | - | 4.8 | 1.13 | 5 | 2.8 | 1.67 | 10.92 | 15.3075 | 48 | 12 | 2 | 50 | 3.6 | N | N | INV | INV | E(2O+D) | OT+ ok; OT- ok | encoder | 0 | 0 | 0 | EncAux | 133.3333 | 10000 | 400 | 2 | -4675 | -175227 | 50653 | 1898173 | 0 | 0 | -175227 | 1898173 |  | 400 | 0.24 | 138.32 | 30000 | 29979.7571 | 30000 | 207.32 | 207.48 | 207.34 | mm | 266.6667 | 266.6667 | 266.67 | 1 | attr://EncEncIn | VALUE/10000 | TRUE | 0 | -17.5227 | 189.8173 | -181.311 | 26.016 | -17.5227 | 189.8173 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2019-03-08 |  | Step_per_unit:266.67;Offset:0;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 207.32 |  | mm | 0
15 | B108A-O02-OPT-BAFF-01 | Baffle Species | VT | mono_baff_vt | Y | N | 2O | W041411 | W041412 |  | - | ok | Oriental PK268M-E2.0 parallel | - | 4.8 | 1.13 | 5 | 2.4 | 1.27 | 9.36 | 15.3075 | 48 | 12 | 2.4 | 50 | 3.1 | N | N | INV | N | 2O | OT+ ok; OT- ok | encoder | 7365 | 0 | 0 | EncAux | 200 | 10000 | 400 | 2 | -8615 | -215224 | 9374 | 232402 | -223630 | 229130 | -215224 | 232402 |  | 600 | 0.2 | 29.9817 | 20000 | 19906.6541 | 20000 | 46mm | 44.9725 | 44.7626 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 0.7365 | 22.2589 | -22.5037 | -22.363 | 22.913 | 22.2589 | -22.5037 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S010321; 348-07-00; | Step_per_unit:400;Offset:0.7365;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 7365 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 46mm |  | mm | 0
16 | B108A-O02-OPT-BAFF-01 | Baffle Species | VB | mono_baff_vb | Y | N | 2O | W041413 | W041414 |  | - | ok | Oriental PK268M-E2.0 parallel | - | 4.8 | 1.13 | 5 | 2.3 | 1.17 | 8.97 | 15.3075 | 48 | 12 | 2.4 | 50 | 3 | N | N | INV | N | 2O | OT+ ok; OT- ok | encoder | 0 | 0 | 0 | EncAux | 200 | 10000 | 400 | 2 | -8790 | -219312 | 9312 | 232728 | -229830 | 229950 | -219312 | 232728 |  | 600 | 0.2 | 30.17 | 20000 | 19977.4611 | 20000 | 46mm | 45.255 | 45.204 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 0 | 21.9312 | -23.2728 | -22.983 | 22.995 | 21.9312 | -23.2728 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S010321; 348-07-00; | Step_per_unit:400;Offset:0;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 46mm |  | mm | 0
17 | B108A-O02-OPT-BAFF-01 | Baffle Species | HR | mono_baff_hr | Y | N | 2O | W041415 | W041416 |  | - | ok | Oriental PK268M-E2.0 parallel | - | 4.8 | 1.13 | 4.9 | 2.3 | 1.17 | 8.97 | 15.0014 | 48 | 12 | 2.4 | 50 | 3 | N | N | INV | N | 2O | OT+ ok; OT- ok | encoder | 3242 | 0 | 0 | EncAux | 200 | 10000 | 400 | 2 | -8845 | -219707 | 9048 | 225898 | -227490 | 228360 | -219707 | 225898 |  | 600 | 0.2 | 29.8217 | 20000 | 19923.0984 | 20000 | 46mm | 44.7325 | 44.5605 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 0.3242 | 22.2949 | -22.2656 | -22.749 | 22.836 | 22.2949 | -22.2656 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S010321; 348-07-00; | Step_per_unit:400;Offset:0.3242;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 3242 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 46mm |  | mm | 0
18 | B108A-O02-OPT-BAFF-01 | Baffle Species | HL | mono_baff_hl | Y | N | 2O | W041417 | W041418 |  | - | ok | Oriental PK268M-E2.0 parallel | - | 4.8 | 1.13 | 5 | 2.3 | 1.17 | 8.97 | 15.3075 | 48 | 12 | 2.4 | 50 | 3.1 | N | N | INV | N | 2O | OT+ ok; OT- ok | encoder | 2744 | 0 | 0 | EncAux | 200 | 10000 | 400 | 2 | -8992 | -223462 | 9134 | 228157 | -230880 | 229170 | -223462 | 228157 |  | 600 | 0.2 | 30.21 | 20000 | 19932.4286 | 20000 | 46mm | 45.315 | 45.1619 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 0.2744 | 22.6206 | -22.5413 | -23.088 | 22.917 | 22.6206 | -22.5413 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S010321; 348-07-00; | Step_per_unit:400;Offset:0.2744;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 2744 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 46mm |  | mm | 0
21 | B108A-O03-OPT-BAFF-01 | Baffle Species | VT | m3_baff_vt | Y | N | GND | W041419 | W041420 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 3.5 | 1.5 | 7 | 2.669 | 48 | 12 | 1 | 50 | 4.2 | N | N | INV | N | GND | N/A | encoder | -11707 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -3674 | -91503 | 4521 | 112692 | -112123 | 88140 | -91503 | 112692 |  | 500 | 0.12 | 16.39 | 10000 | 9966.809 | 10000 | 20mm | 20.4875 | 20.4195 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 1.1707 | 10.321 | -10.0985 | -10.04 | 9.98 | 10.321 | -10.0985 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | W001017/03; | Step_per_unit:400;Offset:1.1707;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -11707 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 20mm |  | mm | 0
22 | B108A-O03-OPT-BAFF-01 | Baffle Species | VB | m3_baff_vb | Y | N | GND | W041421 | W041422 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 3.6 | 1.6 | 7.2 | 2.669 | 48 | 12 | 1 | 50 | 4.2 | N | N | INV | N | GND | N/A | encoder | -6628 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -3862 | -96047 | 4366 | 108841 | -106760 | 93339 | -96047 | 108841 |  | 500 | 0.12 | 16.456 | 10000 | 9960.525 | 10000 | 20mm | 20.57 | 20.4888 | mm | 400 | 400 | 400 | 1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.6628 | -10.2675 | 10.2213 | -10.01 | 10 | -10.2675 | 10.2213 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | W001017/03; enc interface LED red/orange sometimes | Step_per_unit:400;Offset:-0.6628;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -6628 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 20mm |  | mm | 0
23 | B108A-O03-OPT-BAFF-01 | Baffle Species | HL | m3_baff_hl | Y | N | GND | W041423 | W041424 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 3.4 | 1.4 | 6.8 | 2.669 | 48 | 12 | 1 | 50 | 4.2 | N | N | INV | N | GND | N/A | encoder | -1381 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -4164 | -103592 | 4070 | 101328 | -101433 | 98457 | -103592 | 101328 |  | 500 | 0.12 | 16.468 | 10000 | 9954.8215 | 10000 | 20mm | 20.585 | 20.492 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.1381 | 10.2211 | -10.2709 | -10.01 | 9.98 | 10.2211 | -10.2709 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | W001017/03; | Step_per_unit:400;Offset:-0.1381;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -1381 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 20mm |  | mm | 0
24 | B108A-O03-OPT-BAFF-01 | Baffle Species | HR | m3_baff_hr | Y | N | GND | W041425 | W041426 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.8 | 3.8 | 1.8 | 7.6 | 2.826 | 48 | 12 | 1 | 50 | 4.3 | N | N | INV | N | GND | N/A | encoder | -13606 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -4550 | -113304 | 3578 | 89071 | -113511 | 86648 | -113304 | 89071 |  | 500 | 0.12 | 16.256 | 10000 | 9959.3996 | 10000 | 20mm | 20.32 | 20.2375 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | -1.3606 | 9.9698 | -10.2677 | -9.99 | 10.03 | 9.9698 | -10.2677 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | W001017/03; | Step_per_unit:400;Offset:-1.3606;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -13606 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 20mm |  | mm | 0
25 | B108A-O03-OPT-MIR-03 | Mirror Species | VC | m3_vc | Y | N | 2O | W041427 | W041428 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 2.2 | 1.5 | 8.58 | 4.2861 | 48 | 12 | 2 | 30 | 3.1 | N | N | INV | N | 2O |  | absolute | 30254060 | 0 | 32496272 | N/A | 20000 | 1000000 | 400 | 2 | 784714 | 19620285 | 1634575 | 40867859 | 19604250 | 40859680 | 19620285 | 40867859 |  | 5000 | 0.3 | 169.9722 | 20000 | 20000.9875 | 20000 | ±10mm | 21.2465 | 21.2476 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -32.4963 | -12.876 | 8.3716 | ~-10.5 | ~10.5 | -12.876 | 8.3716 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | S010257;691-17-00; | Step_per_unit:40000;Offset:-32.496272;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30254060 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
26 | B108A-O03-OPT-MIR-03 | Mirror Species | VR | m3_vr | Y | N | 2O | W041429 | W041430 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 2.1 | 1.4 | 8.19 | 4.2861 | 48 | 12 | 2 | 30 | 3.1 | N | N | INV | N | 2O |  | absolute | 30657070 | 0 | 32203645 | N/A | 20000 | 1000000 | 400 | 2 | 703801 | 17595466 | 1557026 | 38925716 | 19834590 | 41178010 | 17595466 | 38925716 |  | 5000 | 0.3 | 170.645 | 20000 | 19999.6484 | 20000 | ±10mm | 21.3306 | 21.3302 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -32.2036 | -14.6082 | 6.7221 | ~-10.5 | ~10.5 | -14.6082 | 6.7221 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | S010257;691-17-00; | Step_per_unit:40000;Offset:-32.203645;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30657070 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
27 | B108A-O03-OPT-MIR-03 | Mirror Species | VL | m3_vl | Y | N | 2O | W041431 | W041432 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 2.2 | 1.5 | 8.58 | 4.2861 | 48 | 12 | 2 | 30 | 3 | N | N | INV | N | 2O |  | absolute | 31187290 | 0 | 33689300 | N/A | 20000 | 1000000 | 400 | 2 | 821476 | 20540940 | 1671413 | 41790084 | 20541080 | 40806980 | 20540940 | 41790084 |  | 5000 | 0.3 | 169.9874 | 20000 | 20000.6768 | 20000 | ±10mm | 21.2484 | 21.2491 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -33.6893 | -13.1484 | 8.1008 | ~-10.5 | ~10.5 | -13.1484 | 8.1008 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | S010257;691-17-00; | Step_per_unit:40000;Offset:-33.6893;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 31187290 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
28 | B108A-O03-OPT-MIR-03 | Mirror Species | HLAT | m3_x | Y | N | 2O | W041433 | W041434 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 2.2 | 1.5 | 8.58 | 4.2861 | 48 | 12 | 2 | 30 | 3.1 | N | N | INV | N | 2O |  | absolute | 45348700 | 0 | 45456503 | N/A | 20000 | 1000000 | 400 | 2 | 1388139 | 34708224 | 2240516 | 56015465 | 34695970 | 56029340 | 34708224 | 56015465 |  | 5000 | 0.3 | 170.4754 | 20000 | 19997.9502 | 20000 | ±10mm | 21.3094 | 21.3072 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -45.3487 | -10.6405 | 10.6668 | ~-10.5 | ~10.5 | -10.6405 | 10.6668 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | S010257;691-17-00; | Step_per_unit:40000;Offset:-45.3487;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 45348700 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
31 | B108A-O03-OPT-MIR-03 | Mirror Species | HPIT | m3_hpit | Y | N | 2O | W041435 | W041436 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.39 | 2.2 | 1.5 | 8.58 | 4.2555 | 48 | 12 | 2 | 30 | 3 | N | N | INV | N | 2O |  | absolute | 35612000 | 0 | 34127574 | N/A | 5882.3529 | 294550.81 | 400 | 2 | 1272211 | 31804864 | 1536901 | 38397674 | 32274810 | 38734300 | 31804864 | 38397674 |  | 2000 | 0.3 | 132.345 | 20029.4551 | 19926.1325 | 20000 | ±3mm; ±10 mrad | 22.4987 | 22.3826 | mrad | 11764.7059 | 11782.0324 | 11782 | 1 | attr://EncAbsEnc | VALUE/294550.81 | TRUE | -115.8631 | -7.8856 | 14.497 | ~-3.5 | ~3.5 | -7.8856 | 14.497 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | S010257;691-17-00; | Step_per_unit:11782;Offset:-115.863113735793;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | 0 |  | 35612000 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±3mm; ±10 mrad |  | mrad | 0
32 | B108A-O03-OPT-MIR-03 | Mirror Species | MCHG | m3_mchg | Y | N | 2O | W041437 | W041438 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 3.4 | 1.4 | 6.8 | 2.669 | 48 | 12 | 1 | 30 | 4.7 | N | N | INV | N | 2O |  | encoder | 105974 | 0 | 0 | N/A | 200 | 10000 | 200 | 2 | -2110 | -52628 | 10144 | 253314 | -66549 | 278450 | -52628 | 253314 |  | 250 | 0.3 | 49.016 | 10000 | 9986.6819 | 10000 | ±32mm | 30.635 | 30.5942 | mm | 400 | 400 | 400 | 1 | attr://EncEncIn | VALUE/10000 | TRUE | -10.5974 | -15.8602 | 14.734 | ~-16 | ~16 | -15.8602 | 14.734 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-19 | S010257;691-17-00; enc readhead adjusted | Step_per_unit:400;Offset:-10.5974;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 105974 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±32mm |  | mm | 0
33 | B108A-OA04-DIA-BPM-01 | BPM Species | V | a_bpm | Y | N | GND | W041439 | W041440 |  | - | ok | Wantai 42BYGHM809 | Renishaw Resolute 26 bits BISS-C | 2.8 | 1.8 |  | 2.3 | 0.5 | 5.52 | 0 | 48 | 12 | 2.2 | 50 | 3.1 | N | N | INV | N | GND | N/A | absolute | 0 | 0 | 0 | N/A | 24000 | 20000 | 400 | 2 | 2605139 | 1084326 | 6967888 | 2901526 | 0 | 0 | 1084326 | 2901526 |  | 5000 | 0.5 | 872.5498 | 333.3333 | 333.2211 | 333 | 120mm | 90.8906 | 90.86 | mm | 48000 | 48048.048 | 48000 | 1 | attr://EncAbsEnc | VALUE/20000 | TRUE | 0 | 54.2163 | 145.0763 | 0 | 0 | 54.2163 | 145.0763 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  |  |  | Step_per_unit:48000;Offset:0;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/20000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 120mm |  | mm | 0
34 | B108A-OB04-DIA-BPM-01 | BPM Species | V | b_bpm | Y | N | GND | W041441 | W041442 |  | - | ok | Wantai 42BYGHM809 | Renishaw Resolute 26 bits BISS-C | 2.8 | 1.8 |  |  | - | 0 | 0 | 48 | 12 | 2.2 | 50 | 3.1 | N | N | INV | N | GND | N/A | absolute | 0 | 0 | 0 | N/A | 24000 | 20000 | 400 | 2 | 2438111 | 1013587 | 6719736 | 2797104 | 0 | 0 | 1013587 | 2797104 |  | 5000 | 0.5 | 856.325 | 333.3333 | 333.2411 | 333 | 120mm | 89.2005 | 89.1758 | mm | 48000 | 48048.048 | 48000 | 1 | attr://EncAbsEnc | VALUE/20000 | TRUE | 0 | 50.6793 | 139.8552 | 0 | 0 | 50.6793 | 139.8552 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  |  |  | Step_per_unit:48000;Offset:0;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/20000;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 120mm |  | mm | 0
35 | B108A-OA04-OPT-SLIT-01 | Slits Species | V | a_slit1_v | Y | N | 2O | W041443 | W041444 |  | - | ok | Oriental PK244-01 parallel | - | 3.2 | 1.8 | 1.4 | 3.8 | 2 | 9.12 | 2.6376 | 48 | 12 | 1 | 10 | 4.7 | N | N | INV | N | 2O | OT+ ok; OT- ok | encoder | 0 | 0 | 0 | EncAux | 20 | 20 | 200 | 2 | -40888 | -20445 | 59070 | 29534 | -20448 | 29507 | -20445 | 29534 |  | 5000 | 0.5 | 19.9916 | 200 | 200 | 200 | 4996um; 2498um slit gap | 2498.95 | 2498.95 | um | 40 | 40 | 40 | -1 | attr://EncEncIn | VALUE/20 | TRUE | 1365 | 2387.25 | -111.7 | -110.4 | 2387.4 | 2387.25 | -111.7 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S008833;348-08-00 | Step_per_unit:40;Offset:1365;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/20;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 4996um; 2498um slit gap |  | um | 0
36 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAR | a_slit1_hdar | Y | N | GND | W041445 | #N/A |  | - | ok | McLennan 23HS-108 Mk2 | - | 0.59 | 0.37 | 0.34 | 0.9 | 0.53 | 4.95 | 1.468 | 48 | 12 | 1.2 | 10 | 1.7 | N | N | INV | N | GND | N/A | 0 | 0 | 0 | 0 | LIM+ | 200 | 0 | 200 | 16 | -38007 |  | -39 |  | -18862.4 | 19096.8 | -38007 | -39 |  | 800 | 0.08 | 47.46 | #N/A | #N/A |  | ±12.5mm | 11.865 | #N/A | mm | 3200 | 0 | 3200 | -1 |  |  | FALSE | -5.55 | 6.3272 | -5.5378 | 19.8 | 7.9 | 6.3272 | -5.5378 | - | - | - | - | - | - | DISABLE |  |  | #DIV/0! | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-25 | S008833;348-08-00 | Step_per_unit:3200;Offset:-5.55;Sign:-1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±12.5mm |  | mm | 0
37 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAL | a_slit1_hdal | Y | N | GND | W041446 | #N/A |  | - | ok | McLennan 23HS-108 Mk2 | - | 0.59 | 0.37 | 0.3 | 1 | 0.63 | 5.5 | 1.2953 | 48 | 12 | 1 | 10 | 1.5 | N | N | INV | N | GND | N/A | 0 | 0 | 0 | 0 | LIM- | 200 | 0 | 200 | 16 | 39 |  | 39295 |  | -18467.2 | 19145.6 | 39 | 39295 |  | 800 | 0.08 | 49.07 | #N/A | #N/A |  | ±12.5mm | 12.2675 | #N/A | mm | 3200 | 0 | 3200 | 1 |  |  | FALSE | -5.9 | -5.8878 | 6.3797 | 19.2 | 7.5 | -5.8878 | 6.3797 | - | - | - | - | - | - | DISABLE |  |  | #DIV/0! | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-25 | S008833;348-08-00 | Step_per_unit:3200;Offset:-5.9;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±12.5mm |  | mm | 0
38 | B108A-OA04-OPT-SLIT-01 | Slits Species | HZ | a_slit1_hz | Y | N | 2O | W041447 | W041448 |  | - | ok | Oriental PK268M-E2.0 parallel | - | 4.8 | 1.13 | 4.8 | 1.8 | 0.67 | 7.02 | 14.6952 | 48 | 12 | 2 | 30 | 2.6 | N | N | INV | INV | 2O | OT+ ok; OT- ok | encoder | -0.324 | 0 | 0 | EncAux | 200 | 10000 | 400 | 2 | -19097 | -477983 | 19172 | 479785 | -487370 | 487370 | -477983 | 479785 |  | 1000 | 0.3 | 38.269 | 20000 | 20021.8035 | 20000 | ±50mm | 95.6725 | 95.7768 | mm | 400 | 400 | 400 | 1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.324 | -48.1223 | 47.6545 | ~-45 | ~45 | -48.1223 | 47.6545 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S008833;348-08-00 | Step_per_unit:400;Offset:-0.324;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -0.324 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±50mm |  | mm | 0
0
0
41 | B108A-OB04-OPT-SLIT-01 | Slits Species | V | b_slit1_v | Y | N | 2O | W041449 | W041450 |  | - | ok | Oriental PK244-01 parallel | - | 3.2 | 1.8 | 1.4 | 4 | 2.2 | 9.6 | 2.6376 | 48 | 12 | 1 | 10 | 4.9 | N | N | INV | N | 2O | OT+ ok; OT- ok | encoder | 0 | 0 | 0 | EncAux | 20 | 20 | 200 | 2 | -39437 | -19707 | 58846 | 29377 | -19736 | 29292 | -19707 | 29377 |  | 1000 | 0.3 | 98.283 | 200 | 199.766 | 200 | 4903um; 2451um slit gap | 2457.075 | 2454.2 | um | 40 | 40 | 40 | -1 | attr://EncEncIn | VALUE/20 | TRUE | 1398.3 | 2383.65 | -70.55 | -66.4 | 2385.1 | 2383.65 | -70.55 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S008834; 348-08-00 | Step_per_unit:40;Offset:1398.3;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/20;UseEncoderSource:TRUE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 4903um; 2451um slit gap |  | um | 0
42 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAR | b_slit1_hdar | Y | N | GND | W041451 | #N/A |  | - | ok | McLennan 23HS-108 Mk2 | - | 0.59 | 0.37 | 0.34 | 0.9 | 0.53 | 4.95 | 1.468 | 48 | 12 | 1.2 | 10 | 1.9 | N | N | INV | N | GND | N/A | 0 | 0 | 0 | 0 | LIM+ | 200 | 0 | 200 | 16 | -36904 |  | -8 |  | -19306.4 | 19316.8 | -36904 | -8 |  | 800 | 0.08 | 46.12 | #N/A | #N/A |  | ±12.5mm | 11.53 | #N/A | mm | 3200 | 0 | 3200 | -1 |  |  | FALSE | -5.8 | 5.7325 | -5.7975 | 19.9 | 7.9 | 5.7325 | -5.7975 | - | - | - | - | - | - | DISABLE |  |  | #DIV/0! | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-25 | S008834; 348-08-00 | Step_per_unit:3200;Offset:-5.8;Sign:-1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±12.5mm |  | mm | 0
43 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAL | b_slit1_hdal | Y | N | GND | W041452 | #N/A |  | - | ok | McLennan 23HS-108 Mk2 | - | 0.59 | 0.37 | 0.34 | 0.9 | 0.53 | 4.95 | 1.468 | 48 | 12 | 1 | 10 | 1.6 | N | N | INV | N | GND | N/A | 0 | 0 | 0 | 0 | LIM- | 200 | 0 | 200 | 16 | 25 |  | 35391 |  | -18134.4 | 18988 | 25 | 35391 |  | 800 | 0.08 | 44.2075 | #N/A | #N/A |  | ±12.5mm | 11.0519 | #N/A | mm | 3200 | 0 | 3200 | 1 |  |  | FALSE | -5.55 | -5.5422 | 5.5097 | 18.9 | 7.3 | -5.5422 | 5.5097 | - | - | - | - | - |  | DISABLE |  |  | #DIV/0! | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-10-25 | S008834; 348-08-00 | Step_per_unit:3200;Offset:-5.55;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±12.5mm |  | mm | 0
44 | B108A-OB04-OPT-SLIT-01 | Slits Species | HZ | b_slit1_hz | Y | N | 2O | W041453 | W041454 |  | - | ok | Oriental PK268M-E2.0 parallel | - | 4.8 | 1.13 | 4.6 | 1.8 | 0.67 | 7.02 | 14.0829 | 48 | 12 | 2 | 10 | 2.6 | N | N | INV | INV | 2O | OT+ ok; OT- ok | encoder | -0.2422 | 0 | 0 | EncAux | 200 | 10000 | 400 | 2 | -18973 | -474158 | 19224 | 480557 | -48582 | 478266 | -474158 | 480557 |  | 1000 | 0.3 | 38.197 | 20000 | 19995.6017 | 20000 | ±50mm | 95.4925 | 95.4715 | mm | 400 | 400 | 400 | 1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.2422 | -47.658 | 47.8135 | ~-45 | ~45 | -47.658 | 47.8135 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09 | S008834; 348-08-00 | Step_per_unit:400;Offset:-0.2422;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -0.2422 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±50mm |  | mm | 0
45 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VT | a_m4_baff_vt | Y | N | GND | W041455 | W041456 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.6 | 2.8 | 0.8 | 5.6 | 2.512 | 48 | 12 | 1 | 50 | 3.4 | N | N | INV | N | GND | N/A | encoder | 2065 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -4147 | -102558 | 4184 | 103844 | -98155 | 102132 | -102558 | 103844 |  | 500 | 0.12 | 16.662 | 10000 | 9910.0708 | 10000 | ±10mm | 20.8275 | 20.6402 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 0.2065 | 10.4623 | -10.1779 | -10.02 | 10.01 | 10.4623 | -10.1779 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-14 | W001017/02; | Step_per_unit:400;Offset:0.2065;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 2065 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
46 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VB | a_m4_baff_vb | Y | N | GND | W041457 | W041458 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 2.7 | 0.7 | 5.4 | 2.669 | 48 | 12 | 1 | 50 | 3.4 | N | N | INV | N | GND | N/A | encoder | 7250 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -4049 | -100687 | 4044 | 100578 | -92556 | 107334 | -100687 | 100578 |  | 500 | 0.12 | 16.186 | 10000 | 9947.609 | 10000 | ±10mm | 20.2325 | 20.1265 | mm | 400 | 400 | 400 | 1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.725 | -10.7937 | 9.3328 | -9.98 | 10.01 | -10.7937 | 9.3328 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-14 | W001017/02; | Step_per_unit:400;Offset:-0.725;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 7250 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
47 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HL | a_m4_baff_hl | Y | N | GND | W041459 | W041460 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.6 | 2.8 | 0.8 | 5.6 | 2.512 | 48 | 12 | 1 | 50 | 3.4 | N | N | INV | N | GND | N/A | encoder | -15490 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -3442 | -85850 | 4683 | 116528 | -115478 | 84576 | -85850 | 116528 |  | 500 | 0.12 | 16.25 | 10000 | 9963.2246 | 10000 | ±10mm | 20.3125 | 20.2378 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 1.549 | 10.134 | -10.1038 | -10 | 10.01 | 10.134 | -10.1038 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-14 | W001017/02; | Step_per_unit:400;Offset:1.549;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -15490 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
48 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HR | a_m4_baff_hr | Y | N | GND | W041461 | W041462 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 2.8 | 0.8 | 5.6 | 2.669 | 48 | 12 | 1 | 50 | 3.4 | N | N | INV | N | GND | N/A | encoder | 13248 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -3286 | -82027 | 4715 | 117421 | -86695 | 112747 | -82027 | 117421 |  | 500 | 0.12 | 16.002 | 10000 | 9971.1536 | 10000 | ±10mm | 20.0025 | 19.9448 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 1.3248 | 9.5275 | -10.4173 | -9.99 | 9.95 | 9.5275 | -10.4173 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-14 | W001017/02; | Step_per_unit:400;Offset:1.3248;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 13248 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
0
0
51 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VT | b_m4_baff_vt | Y | N | GND | W041463 | W041464 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 2.8 | 0.8 | 5.6 | 2.669 | 48 | 12 | 1 | 50 | 3.5 | N | N | INV | N | GND | N/A | encoder | 8984 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -4344 | -107986 | 3572 | 89082 | -90968 | 108995 | -107986 | 89082 |  | 500 | 0.12 | 15.832 | 10000 | 9957.9586 | 10000 | ±10mm | 19.79 | 19.7068 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.8984 | 9.9002 | -9.8066 | -10 | 10 | 9.9002 | -9.8066 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-25 | W001017/01; broken motor connector was replaced | Step_per_unit:400;Offset:-0.8984;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 8984 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
52 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VB | b_m4_baff_vb | Y | N | GND | W041465 | W041466 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 2.8 | 0.8 | 5.6 | 2.669 | 48 | 12 | 1 | 50 | 3.7 | N | N | INV | N | GND | N/A | encoder | -3793 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -3740 | -93171 | 4178 | 104074 | -103536 | 96474 | -93171 | 104074 |  | 500 | 0.12 | 15.836 | 10000 | 9964.3849 | 10000 | ±10mm | 19.795 | 19.7245 | mm | 400 | 400 | 400 | 1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.3793 | -9.6964 | 10.0281 | -9.97 | 10.03 | -9.6964 | 10.0281 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-25 | W001017/01; | Step_per_unit:400;Offset:-0.3793;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | -3793 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
53 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HL | b_m4_baff_hl | Y | N | GND | W041467 | W041468 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 2.8 | 0.8 | 5.6 | 2.669 | 48 | 12 | 1 | 50 | 3.5 | N | N | INV | N | GND | N/A | encoder | 6826 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -4229 | -105181 | 3798 | 94569 | -93415 | 106568 | -105181 | 94569 |  | 500 | 0.12 | 16.054 | 10000 | 9953.9056 | 10000 | ±10mm | 20.0675 | 19.975 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | -0.6826 | 9.8355 | -10.1395 | -10.02 | 9.97 | 9.8355 | -10.1395 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-25 | W001017/01; | Step_per_unit:400;Offset:-0.6826;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 6826 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
54 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HR | b_m4_baff_hr | Y | N | GND | W041469 | W041470 |  | - | ok | Nanotec L4118S1404-M6X1 | - | 2.8 | 2 | 1.7 | 2.8 | 0.8 | 5.6 | 2.669 | 48 | 12 | 1 | 50 | 3.7 | N | N | INV | N | GND | N/A | encoder | 9685 | 0 | 0 | EncAux | 200 | 10000 | 200 | 2 | -3814 | -95165 | 4251 | 105619 | -90340 | 109821 | -95165 | 105619 |  | 500 | 0.12 | 16.13 | 10000 | 9958.2889 | 10000 | ±10mm | 20.1625 | 20.0784 | mm | 400 | 400 | 400 | -1 | attr://EncEncIn | VALUE/10000 | TRUE | 0.9685 | 10.485 | -9.5934 | -10 | 10.01 | 10.485 | -9.5934 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 10um | 5um |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-25 | W001017/01; | Step_per_unit:400;Offset:0.9685;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | 0 |  | 9685 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
55 | B108A-OA05-OPT-MIR-04 | Mirror Species | VC | a_m4_vc | Y | N | 2O | W041471 | W041472 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.4 | 0.7 | 5.46 | 4.2861 | 48 | 12 | 2.5 | 50 | 2 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30511950 | 0 | 30912411 | N/A | 20000 | 1000000 | 400 | 2 | 795297 | 19884048 | 1645900 | 41151584 | 19864150 | 41147710 | 19884048 | 41151584 |  | 5000 | 0.3 | 170.1206 | 20000 | 20002.3146 | 20000 | ±10mm | 21.2651 | 21.2675 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -31.9534 | -12.0693 | 9.1982 | ~-10.5 | ~10.5 | -12.0693 | 9.1982 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-26 | S010303; 691-18-00; | Step_per_unit:40000;Offset:-31.953371;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30511950 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
56 | B108A-OA05-OPT-MIR-04 | Mirror Species | VR | a_m4_vr | Y | N | 2O | W041473 | W041474 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.4 | 0.7 | 5.46 | 4.2861 | 48 | 12 | 2.5 | 50 | 2.1 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30028150 | 0 | 29866607 | N/A | 20000 | 1000000 | 400 | 2 | 780642 | 19516365 | 1625100 | 40630141 | 19441640 | 40571580 | 19516365 | 40630141 |  | 5000 | 0.3 | 168.8916 | 20000 | 20002.2035 | 20000 | ±10mm | 21.1115 | 21.1138 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -32.0109 | -12.4945 | 8.6193 | ~-10.5 | ~10.5 | -12.4945 | 8.6193 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-26 | S010303; 691-18-00; | Step_per_unit:40000;Offset:-32.01086;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30028150 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
57 | B108A-OA05-OPT-MIR-04 | Mirror Species | VL | a_m4_vl | Y | N | 2O | W041475 | W041476 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.3 | 0.6 | 5.07 | 4.2861 | 48 | 12 | 2.5 | 50 | 2 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30551700 | 0 | 30232634 | N/A | 20000 | 1000000 | 400 | 2 | 794924 | 19869761 | 1643101 | 41080509 | 19848610 | 41080910 | 19869761 | 41080509 |  | 5000 | 0.3 | 169.6354 | 20000 | 20005.9638 | 20000 | ±10mm | 21.2044 | 21.2107 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -31.4129 | -11.5431 | 9.6676 | ~-10.5 | ~10.5 | -11.5431 | 9.6676 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-26 | S010303; 691-18-00; | Step_per_unit:40000;Offset:-31.412874;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30551700 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
58 | B108A-OA05-OPT-MIR-04 | Mirror Species | HLAT | a_m4_x | Y | N | 2O | W041477 | W041478 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.3 | 0.6 | 5.07 | 4.2861 | 48 | 12 | 2 | 30 | 2 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 45038630 | 0 | 42320617 | N/A | 20000 | 1000000 | 400 | 2 | 1376319 | 34415516 | 2232252 | 55807958 | 34402140 | 55824330 | 34415516 | 55807958 |  | 5000 | 0.3 | 171.1866 | 20000 | 19994.5014 | 20000 | ±10mm | 21.3983 | 21.3924 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -45.0386 | -10.6231 | 10.7693 | ~-10.5 | ~10.5 | -10.6231 | 10.7693 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-27 | S010303; 691-18-00; | Step_per_unit:40000;Offset:-45.03863;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 45038630 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
0
0
61 | B108A-OA05-OPT-MIR-04 | Mirror Species | HPIT | a_m4_hpit | Y | N | 2O | W041479 | W041480 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.3 | 0.6 | 5.07 | 4.2861 | 48 | 12 | 2 | 30 | 2 | N | N | INV | N | 2O | OT+ nok; OT-nok | absolute | 35837070 | 0 | 36806540 | N/A | 5882.3529 | 294550.81 | 400 | 2 | 1311544 | 32791845 | 1557582 | 38939784 | 32778150 | 38942390 | 32791845 | 38939784 |  | 2000 | 0.3 | 123.019 | 20029.4551 | 19990.2096 | 20000 | ±3mm; ±10 mrad | 20.9132 | 20.8723 | mrad | 11764.7059 | 11782.0324 | 11782 | 1 | attr://EncAbsEnc | VALUE/294550.81 | TRUE | -125.6299 | -14.3016 | 6.5707 | ~-3.5 | ~3.5 | -14.3016 | 6.5707 | - | - | - | - | - | - | NONE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-27 | S010303; 691-18-00; 2019-03-20 overtravel not working; when presses it jumps back but later is activated again | Step_per_unit:11782;Offset:-125.629893871281;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | 0 |  | 35837070 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±3mm; ±10 mrad |  | mrad | 0
62 | B108A-OB05-OPT-MIR-04 | Mirror Species | VC | b_m4_vc | Y | N | 2O | W041481 | W041482 |  | - | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.7 | 1 | 6.63 | 4.2861 | 48 | 12 | 2.5 | 50 | 2.4 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30239670 | 0 | 30225090 | N/A | 20000 | 1000000 | 400 | 2 | 784765 | 19619107 | 1636727 | 40919850 | 19580510 | 40900370 | 19619107 | 40919850 |  | 5000 | 0.3 | 170.3924 | 20000 | 20001.5897 | 20000 | ±10mm | 21.2991 | 21.3007 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -32.2327 | -12.6136 | 8.6871 | ~-10.5 | ~10.5 | -12.6136 | 8.6871 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-26 | S010304; 691-19-00; | Step_per_unit:40000;Offset:-32.232744;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30239670 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
63 | B108A-OB05-OPT-MIR-04 | Mirror Species | VR | b_m4_vr | Y | N | 2O | W041483 | W041484 |  | 0 | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.8 | 1.1 | 7.02 | 4.2861 | 48 | 12 | 2.5 | 50 | 2.4 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30361730 | 0 | 31022503 | N/A | 20000 | 1000000 | 400 | 2 | 788398 | 19714685 | 1636972 | 40926303 | 19719320 | 40947080 | 19714685 | 40926303 |  | 5000 | 0.3 | 169.7148 | 20000 | 19997.4244 | 20000 | ±10mm | 21.2143 | 21.2116 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -32.1789 | -12.4642 | 8.7474 | ~-10.5 | ~10.5 | -12.4642 | 8.7474 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-26 | S010304; 691-19-00; | Step_per_unit:40000;Offset:-32.17886;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30361730 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
64 | B108A-OB05-OPT-MIR-04 | Mirror Species | VL | b_m4_vl | Y | N | 2O | W041485 | W041486 |  | 0 | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.7 | 1 | 6.63 | 4.2861 | 48 | 12 | 2.5 | 50 | 2.4 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 30286210 | 0 | 28388395 | N/A | 20000 | 1000000 | 400 | 2 | 786089 | 19651803 | 1631892 | 40798957 | 19625100 | 40805940 | 19651803 | 40798957 |  | 5000 | 0.3 | 169.1606 | 20000 | 20001.9664 | 20000 | ±10mm | 21.1451 | 21.1472 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -31.5288 | -11.877 | 9.2701 | ~-10.5 | ~10.5 | -11.877 | 9.2701 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-26 | S010304; 691-19-00; | Step_per_unit:40000;Offset:-31.528825;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 30286210 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
65 | B108A-OB05-OPT-MIR-04 | Mirror Species | HLAT | b_m4_x | Y | N | 2O | W041487 | W041488 |  | 0 | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.5 | 1.8 | 1.1 | 7.02 | 4.5922 | 48 | 12 | 2 | 50 | 2.4 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 45573530 | 0 | 43744808 | N/A | 20000 | 1000000 | 400 | 2 | 1398438 | 34989446 | 2245617 | 56148575 | 34978650 | 56166560 | 34989446 | 56148575 |  | 5000 | 0.3 | 169.4358 | 20000 | 19980.7871 | 20000 | ±10mm | 21.1795 | 21.1591 | mm | 40000 | 40000 | 40000 | 1 | attr://EncAbsEnc | VALUE/1000000 | TRUE | -45.5735 | -10.5841 | 10.575 | ~-10.5 | ~10.5 | -10.5841 | 10.575 | - | - | - | - | - |  | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-25 | S010304; 691-19-00; | Step_per_unit:40000;Offset:-45.57353;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | 0 |  | 45573530 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±10mm |  | mm | 0
66 | B108A-OB05-OPT-MIR-04 | Mirror Species | HPIT | b_m4_hpit | Y | N | 2O | W041489 | W041490 |  | 0 | ok | Oriental PK264M-E2.0 parallel | Renishaw Resolute 32 bits BISS-C | 1.7 | 0.7 | 1.4 | 1.8 | 1.1 | 7.02 | 4.2861 | 48 | 12 | 2 | 50 | 2.4 | N | N | INV | N | 2O | OT+ ok; OT- ok | absolute | 35812410 | 0 | 35697136 | N/A | 5882.3529 | 294550.81 | 400 | 2 | 1308049 | 32701460 | 1558028 | 38944653 | 32694310 | 38960370 | 32701460 | 38944653 |  | 2000 | 0.3 | 124.9895 | 20029.4551 | 19979.8959 | 20000 | ±3mm; ±10 mrad | 21.2482 | 21.1956 | mrad | 11764.7059 | 11782.0324 | 11782 | 1 | attr://EncAbsEnc | VALUE/294550.81 | TRUE | -119.7453 | -8.7238 | 12.4718 | ~-3.5 | ~3.5 | -8.7238 | 12.4718 | - | - | - | - | - | - | DISABLE |  |  | 0 | 0 | 0 | 0 |  |  |  |  | 0 |  | Gabriela Todorescu; 2017-09-25 | S010304; 691-19-00; | Step_per_unit:11782;Offset:-119.745275865987;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | 0 |  | 35812410 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | ±3mm; ±10 mrad |  | mrad | 0
67 | B108A-FE-DIA-SCRN-01 | Fluorescence Screen Front End 1.5GeV | Y | 0 | Y | N | 0 | #N/A | #N/A |  | 0 |  | 0 | 0 | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | [9 empty] | 0 |  | 0 | 0 | 0 | 0 |  | 0 | 0 | #N/A | 2 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! | #N/A | #N/A |  | 0 |  | #N/A | mm | 0 | 0 | 1 | 1 |  |  |  | 0 | #DIV/0! | #DIV/0! | 0 | 0 |  |  | - | - | - | - | - | - | DISABLE |  |  | #DIV/0! | #DIV/0! | 0 | 0 |  |  |  |  | 0 |  |  |  | Step_per_unit:1;Offset:0;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 0 |  | mm | 0
68 | 0 | 0 | 0 | 0 | Y | N | 0 | #N/A | #N/A |  | #N/A |  | 0 | 0 | #N/A | #N/A |  |  | #N/A | #N/A | #N/A | [9 empty] | 0 |  | 0 | 0 | 0 | 0 |  | 0 | 0 | #N/A | 0 |  |  |  |  | 0 | 0 | 0 | 0 |  |  |  | #DIV/0! | #N/A | #N/A |  | 0 |  | #N/A | 0 | 0 | 0 | 1 | 1 |  |  |  | 0 | #DIV/0! | #DIV/0! | 0 | 0 |  |  | - | - | - | - | - | - |  |  |  | #DIV/0! | #DIV/0! | 0 | 0 |  |  |  |  | 0 |  |  |  | Step_per_unit:1;Offset:0;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | 0 |  | 0 | 0 |  | 0 | 0 |  |  | #DIV/0! |  |  | 0 |  | 0 | 0
0
0
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [53 empty] | 2065 | [19 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -
[5 empty] | 0 | 0 | [73 empty] | -

## Tango
axis | Equipment code | Equiment name | Axis name | Controller dns name | controller | Instrument | devicename | name | Steps per unit | attributes | unit | Soft lim- (units) | Soft lim+ (units) | Ready. (N or blank: don't take the data here)
1 | B108A-FE-OPT-MM-01 | Movable mask unit | X | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /FE | b108a-fe/opt/mm-01-x | mm1_x | 400 | Step_per_unit:400;Offset:48.161702;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -5.5433 | 5.4235
2 | B108A-FE-OPT-MM-01 | Movable mask unit | Y | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /FE | b108a-fe/opt/mm-01-y | mm1_y | 1600 | Step_per_unit:1600;Offset:48.410419;Sign:-1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -5.117 | 5.0962
3 | B108A-FE-OPT-MM-02 | Movable mask unit | X | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /FE | b108a-fe/opt/mm-02-x | mm2_x | 400 | Step_per_unit:400;Offset:-48.999867;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -5.3481 | 5.5431
4 | B108A-FE-OPT-MM-02 | Movable mask unit | Y | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /FE | b108a-fe/opt/mm-02-y | mm2_y | 1600 | Step_per_unit:1600;Offset:-48.306431;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm | -4.9243 | 5.0882
5 | B108A-O01-OPT-MIR-01 | Mirror Species | VC | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O01 | b108a-o01/opt/mir-01-vc | m1_vc | 40000 | Step_per_unit:40000;Offset:-32.419926;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
6 | B108A-O01-OPT-MIR-01 | Mirror Species | VR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O01 | b108a-o01/opt/mir-01-vr | m1_vr | 40000 | Step_per_unit:40000;Offset:-31.990461;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
7 | B108A-O01-OPT-MIR-01 | Mirror Species | VL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O01 | b108a-o01/opt/mir-01-vl | m1_vl | 40000 | Step_per_unit:40000;Offset:-33.327755;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
8 | B108A-O01-OPT-MIR-01 | Mirror Species | HLAT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O01 | b108a-o01/opt/mir-01-hlat | m1_x | 40000 | Step_per_unit:40000;Offset:-45.46146;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
0 | 0 | 0 | 0 |  |  |  | motor//0 | 0 | 0 | 0 | 0
11 | B108A-O01-OPT-MIR-01 | Mirror Species | HPIT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O01 | b108a-o01/opt/mir-01-hpit | m1_hpit | 11782 | Step_per_unit:11782;Offset:-119.547282860977;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | mrad
12 | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PIT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O02 | b108a-o02/opt/mono-01-m2pit | mono_m2pit | 30879.44 | Step_per_unit:30879.44;Offset:0;Sign:1;EncoderSource:tango://b-v-species-csdb-0:10000/B108A/CTL/IK220-02/AnglesMean;EncoderSourceFormula:((VALUE*3.141559265)/180)*1000;UseEncoderSource:TRUE | mrad
13 | B108A-O02-OPT-MONO-01 | Monochromator Species | GPIT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O02 | b108a-o02/opt/mono-01-gpit | mono_gpit | 30879.44 | Step_per_unit:30879.44;Offset:0;Sign:1;EncoderSource:tango://b-v-species-csdb-0:10000/B108A/CTL/IK220-01/AnglesMean;EncoderSourceFormula:((VALUE*3.141559265)/180)*1000;UseEncoderSource:TRUE | mrad
14 | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXC | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O02 | b108a-o02/opt/mono-01-gexc | mono_gexc | 266.67 | Step_per_unit:266.67;Offset:0;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
15 | B108A-O02-OPT-BAFF-01 | Baffle Species | VT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O02 | b108a-o02/opt/baff-01-vt | mono_baff_vt | 400 | Step_per_unit:400;Offset:0.7365;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
16 | B108A-O02-OPT-BAFF-01 | Baffle Species | VB | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O02 | b108a-o02/opt/baff-01-vb | mono_baff_vb | 400 | Step_per_unit:400;Offset:0;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
17 | B108A-O02-OPT-BAFF-01 | Baffle Species | HR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O02 | b108a-o02/opt/baff-01-hr | mono_baff_hr | 400 | Step_per_unit:400;Offset:0.3242;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
18 | B108A-O02-OPT-BAFF-01 | Baffle Species | HL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O02 | b108a-o02/opt/baff-01-hl | mono_baff_hl | 400 | Step_per_unit:400;Offset:0.2744;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
0 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0
21 | B108A-O03-OPT-BAFF-01 | Baffle Species | VT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/baff-01-vt | m3_baff_vt | 400 | Step_per_unit:400;Offset:1.1707;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
22 | B108A-O03-OPT-BAFF-01 | Baffle Species | VB | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/baff-01-vb | m3_baff_vb | 400 | Step_per_unit:400;Offset:-0.6628;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
23 | B108A-O03-OPT-BAFF-01 | Baffle Species | HL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/baff-01-hl | m3_baff_hl | 400 | Step_per_unit:400;Offset:-0.1381;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
24 | B108A-O03-OPT-BAFF-01 | Baffle Species | HR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/baff-01-hr | m3_baff_hr | 400 | Step_per_unit:400;Offset:-1.3606;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
25 | B108A-O03-OPT-MIR-03 | Mirror Species | VC | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/mir-03-vc | m3_vc | 40000 | Step_per_unit:40000;Offset:-32.496272;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
26 | B108A-O03-OPT-MIR-03 | Mirror Species | VR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/mir-03-vr | m3_vr | 40000 | Step_per_unit:40000;Offset:-32.203645;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
27 | B108A-O03-OPT-MIR-03 | Mirror Species | VL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/mir-03-vl | m3_vl | 40000 | Step_per_unit:40000;Offset:-33.6893;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
28 | B108A-O03-OPT-MIR-03 | Mirror Species | HLAT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/mir-03-hlat | m3_x | 40000 | Step_per_unit:40000;Offset:-45.3487;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
0 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0
0 | 0 | 0 | 0 |  |  |  |  | 0 | 0 | 0 | 0
31 | B108A-O03-OPT-MIR-03 | Mirror Species | HPIT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/mir-03-hpit | m3_hpit | 11782 | Step_per_unit:11782;Offset:-115.863113735793;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | mrad
32 | B108A-O03-OPT-MIR-03 | Mirror Species | MCHG | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /O03 | b108a-o03/opt/mir-03-mchg | m3_mchg | 400 | Step_per_unit:400;Offset:-10.5974;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
33 | B108A-OA04-DIA-BPM-01 | BPM Species | V | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA04 | b108a-oa04/dia/bpm-01-v | a_bpm | 48000 | Step_per_unit:48000;Offset:0;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/20000;UseEncoderSource:TRUE | mm
34 | B108A-OB04-DIA-BPM-01 | BPM Species | V | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB04 | b108a-ob04/dia/bpm-01-v | b_bpm | 48000 | Step_per_unit:48000;Offset:0;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/20000;UseEncoderSource:TRUE | mm
35 | B108A-OA04-OPT-SLIT-01 | Slits Species | V | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA04 | b108a-oa04/opt/slit-01-v | a_slit1_v | 40 | Step_per_unit:40;Offset:1365;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/20;UseEncoderSource:TRUE | um
36 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA04 | b108a-oa04/opt/slit-01-hdar | a_slit1_hdar | 3200 | Step_per_unit:3200;Offset:-5.55;Sign:-1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | mm
37 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDAL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA04 | b108a-oa04/opt/slit-01-hdal | a_slit1_hdal | 3200 | Step_per_unit:3200;Offset:-5.9;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | mm
38 | B108A-OA04-OPT-SLIT-01 | Slits Species | HZ | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA04 | b108a-oa04/opt/slit-01-hz | a_slit1_hz | 400 | Step_per_unit:400;Offset:-0.324;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
41 | B108A-OB04-OPT-SLIT-01 | Slits Species | V | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB04 | b108a-ob04/opt/slit-01-v | b_slit1_v | 40 | Step_per_unit:40;Offset:1398.3;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/20;UseEncoderSource:TRUE | um
42 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB04 | b108a-ob04/opt/slit-01-hdar | b_slit1_hdar | 3200 | Step_per_unit:3200;Offset:-5.8;Sign:-1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | mm
43 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDAL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB04 | b108a-ob04/opt/slit-01-hdal | b_slit1_hdal | 3200 | Step_per_unit:3200;Offset:-5.55;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource:FALSE | mm
44 | B108A-OB04-OPT-SLIT-01 | Slits Species | HZ | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB04 | b108a-ob04/opt/slit-01-hz | b_slit1_hz | 400 | Step_per_unit:400;Offset:-0.2422;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
45 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/baff-01-vt | a_m4_baff_vt | 400 | Step_per_unit:400;Offset:0.2065;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
46 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VB | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/baff-01-vb | a_m4_baff_vb | 400 | Step_per_unit:400;Offset:-0.725;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
47 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/baff-01-hl | a_m4_baff_hl | 400 | Step_per_unit:400;Offset:1.549;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
48 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/baff-01-hr | a_m4_baff_hr | 400 | Step_per_unit:400;Offset:1.3248;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
51 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/baff-01-vt | b_m4_baff_vt | 400 | Step_per_unit:400;Offset:-0.8984;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
52 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VB | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/baff-01-vb | b_m4_baff_vb | 400 | Step_per_unit:400;Offset:-0.3793;Sign:1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
53 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/baff-01-hl | b_m4_baff_hl | 400 | Step_per_unit:400;Offset:-0.6826;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
54 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/baff-01-hr | b_m4_baff_hr | 400 | Step_per_unit:400;Offset:0.9685;Sign:-1;EncoderSource:attr://EncEncIn;EncoderSourceFormula:VALUE/10000;UseEncoderSource:TRUE | mm
55 | B108A-OA05-OPT-MIR-04 | Mirror Species | VC | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/mir-04-vc | a_m4_vc | 40000 | Step_per_unit:40000;Offset:-31.953371;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
56 | B108A-OA05-OPT-MIR-04 | Mirror Species | VR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/mir-04-vr | a_m4_vr | 40000 | Step_per_unit:40000;Offset:-32.01086;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
57 | B108A-OA05-OPT-MIR-04 | Mirror Species | VL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/mir-04-vl | a_m4_vl | 40000 | Step_per_unit:40000;Offset:-31.412874;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
58 | B108A-OA05-OPT-MIR-04 | Mirror Species | HLAT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/mir-04-hlat | a_m4_x | 40000 | Step_per_unit:40000;Offset:-45.03863;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
61 | B108A-OA05-OPT-MIR-04 | Mirror Species | HPIT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OA05 | b108a-oa05/opt/mir-04-hpit | a_m4_hpit | 11782 | Step_per_unit:11782;Offset:-125.629893871281;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | mrad
62 | B108A-OB05-OPT-MIR-04 | Mirror Species | VC | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/mir-04-vc | b_m4_vc | 40000 | Step_per_unit:40000;Offset:-32.232744;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
63 | B108A-OB05-OPT-MIR-04 | Mirror Species | VR | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/mir-04-vr | b_m4_vr | 40000 | Step_per_unit:40000;Offset:-32.17886;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
64 | B108A-OB05-OPT-MIR-04 | Mirror Species | VL | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/mir-04-vl | b_m4_vl | 40000 | Step_per_unit:40000;Offset:-31.528825;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
65 | B108A-OB05-OPT-MIR-04 | Mirror Species | HLAT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/mir-04-hlat | b_m4_x | 40000 | Step_per_unit:40000;Offset:-45.57353;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/1000000;UseEncoderSource:TRUE | mm
66 | B108A-OB05-OPT-MIR-04 | Mirror Species | HPIT | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl | /OB05 | b108a-ob05/opt/mir-04-hpit | b_m4_hpit | 11782 | Step_per_unit:11782;Offset:-119.745275865987;Sign:1;EncoderSource:attr://EncAbsEnc;EncoderSourceFormula:VALUE/294550.81;UseEncoderSource:TRUE | mrad
67 | B108A-FE-DIA-SCRN-01 | Fluorescence Screen Front End 1.5GeV | Y | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl |  |  | 0 | 1 | Step_per_unit:1;Offset:0;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | mm
68 | 0 | 0 | 0 | b108a-d100830-cab04-ctl-ipap-01 | icepap_ctrl |  |  | 0 | 1 | Step_per_unit:1;Offset:0;Sign:1;EncoderSource:;EncoderSourceFormula:;UseEncoderSource: | 0

## MOTION1
Cable ID | Conf. Code | Equip. A | Type A | Chan. A | Equip. B | Type B | Chan. B | Chan. B cleaned | AUX
W041387 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML1 | B108A-FE-OPT-MM-01 | Movable mask unit | XML | X | B108A-FE-OPT-MM-01X
W041389 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML2 | B108A-FE-OPT-MM-01 | Movable mask unit | YML | Y | B108A-FE-OPT-MM-01Y
W041391 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML3 | B108A-FE-OPT-MM-02 | Movable mask unit | XML | X | B108A-FE-OPT-MM-02X
W041393 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML4 | B108A-FE-OPT-MM-02 | Movable mask unit | YML | Y | B108A-FE-OPT-MM-02Y
W041395 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML5 | B108A-O01-OPT-MIR-01 | Mirror Species | VCML | VC | B108A-O01-OPT-MIR-01VC
W041397 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML6 | B108A-O01-OPT-MIR-01 | Mirror Species | VRML | VR | B108A-O01-OPT-MIR-01VR
W041399 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML7 | B108A-O01-OPT-MIR-01 | Mirror Species | VLML | VL | B108A-O01-OPT-MIR-01VL
W041401 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | ML8 | B108A-O01-OPT-MIR-01 | Mirror Species | HLATML | HLAT | B108A-O01-OPT-MIR-01HLAT
W041403 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML1 | B108A-O01-OPT-MIR-01 | Mirror Species | HPITML | HPIT | B108A-O01-OPT-MIR-01HPIT
W041405 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML2 | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PITML | M2PIT | B108A-O02-OPT-MONO-01M2PIT
W041407 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML3 | B108A-O02-OPT-MONO-01 | Monochromator Species | GPITML | GPIT | B108A-O02-OPT-MONO-01GPIT
W041409 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML4 | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXCML | GEXC | B108A-O02-OPT-MONO-01GEXC
W041411 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML5 | B108A-O02-OPT-BAFF-01 | Baffle Species | VTML | VT | B108A-O02-OPT-BAFF-01VT
W041413 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML6 | B108A-O02-OPT-BAFF-01 | Baffle Species | VBML | VB | B108A-O02-OPT-BAFF-01VB
W041415 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML7 | B108A-O02-OPT-BAFF-01 | Baffle Species | HRML | HR | B108A-O02-OPT-BAFF-01HR
W041417 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | ML8 | B108A-O02-OPT-BAFF-01 | Baffle Species | HLML | HL | B108A-O02-OPT-BAFF-01HL
W041419 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML1 | B108A-O03-OPT-BAFF-01 | Baffle Species | VTML | VT | B108A-O03-OPT-BAFF-01VT
W041421 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML2 | B108A-O03-OPT-BAFF-01 | Baffle Species | VBML | VB | B108A-O03-OPT-BAFF-01VB
W041423 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML3 | B108A-O03-OPT-BAFF-01 | Baffle Species | HLML | HL | B108A-O03-OPT-BAFF-01HL
W041425 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML4 | B108A-O03-OPT-BAFF-01 | Baffle Species | HRML | HR | B108A-O03-OPT-BAFF-01HR
W041427 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML5 | B108A-O03-OPT-MIR-03 | Mirror Species | VCML | VC | B108A-O03-OPT-MIR-03VC
W041429 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML6 | B108A-O03-OPT-MIR-03 | Mirror Species | VRML | VR | B108A-O03-OPT-MIR-03VR
W041431 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML7 | B108A-O03-OPT-MIR-03 | Mirror Species | VLML | VL | B108A-O03-OPT-MIR-03VL
W041433 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | ML8 | B108A-O03-OPT-MIR-03 | Mirror Species | HLATML | HLAT | B108A-O03-OPT-MIR-03HLAT
W041435 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML1 | B108A-O03-OPT-MIR-03 | Mirror Species | HPITML | HPIT | B108A-O03-OPT-MIR-03HPIT
W041437 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML2 | B108A-O03-OPT-MIR-03 | Mirror Species | MCHGML | MCHG | B108A-O03-OPT-MIR-03MCHG
W041439 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML3 | B108A-OA04-DIA-BPM-01 | BPM Species | VML | V | B108A-OA04-DIA-BPM-01V
W041441 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML4 | B108A-OB04-DIA-BPM-01 | BPM Species | VML | V | B108A-OB04-DIA-BPM-01V
W041443 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML5 | B108A-OA04-OPT-SLIT-01 | Slits Species | VML | V | B108A-OA04-OPT-SLIT-01V
W041445 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML6 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDARML | HDAR | B108A-OA04-OPT-SLIT-01HDAR
W041446 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML7 | B108A-OA04-OPT-SLIT-01 | Slits Species | HDALML | HDAL | B108A-OA04-OPT-SLIT-01HDAL
W041447 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | ML8 | B108A-OA04-OPT-SLIT-01 | Slits Species | HZML | HZ | B108A-OA04-OPT-SLIT-01HZ
W041449 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML1 | B108A-OB04-OPT-SLIT-01 | Slits Species | VML | V | B108A-OB04-OPT-SLIT-01V
W041451 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML2 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDARML | HDAR | B108A-OB04-OPT-SLIT-01HDAR
W041452 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML3 | B108A-OB04-OPT-SLIT-01 | Slits Species | HDALML | HDAL | B108A-OB04-OPT-SLIT-01HDAL
W041453 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML4 | B108A-OB04-OPT-SLIT-01 | Slits Species | HZML | HZ | B108A-OB04-OPT-SLIT-01HZ
W041455 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML5 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VTML | VT | B108A-OA05-OPT-BAFF-01VT
W041457 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML6 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VBML | VB | B108A-OA05-OPT-BAFF-01VB
W041459 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML7 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HLML | HL | B108A-OA05-OPT-BAFF-01HL
W041461 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | ML8 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HRML | HR | B108A-OA05-OPT-BAFF-01HR
W041463 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML1 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VTML | VT | B108A-OB05-OPT-BAFF-01VT
W041465 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML2 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VBML | VB | B108A-OB05-OPT-BAFF-01VB
W041467 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML3 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HLML | HL | B108A-OB05-OPT-BAFF-01HL
W041469 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML4 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HRML | HR | B108A-OB05-OPT-BAFF-01HR
W041471 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML5 | B108A-OA05-OPT-MIR-04 | Mirror Species | VCML | VC | B108A-OA05-OPT-MIR-04VC
W041473 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML6 | B108A-OA05-OPT-MIR-04 | Mirror Species | VRML | VR | B108A-OA05-OPT-MIR-04VR
W041475 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML7 | B108A-OA05-OPT-MIR-04 | Mirror Species | VLML | VL | B108A-OA05-OPT-MIR-04VL
W041477 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | ML8 | B108A-OA05-OPT-MIR-04 | Mirror Species | HLATML | HLAT | B108A-OA05-OPT-MIR-04HLAT
W041479 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML1 | B108A-OA05-OPT-MIR-04 | Mirror Species | HPITML | HPIT | B108A-OA05-OPT-MIR-04HPIT
W041481 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML2 | B108A-OB05-OPT-MIR-04 | Mirror Species | VCML | VC | B108A-OB05-OPT-MIR-04VC
W041483 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML3 | B108A-OB05-OPT-MIR-04 | Mirror Species | VRML | VR | B108A-OB05-OPT-MIR-04VR
W041485 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML4 | B108A-OB05-OPT-MIR-04 | Mirror Species | VLML | VL | B108A-OB05-OPT-MIR-04VL
W041487 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML5 | B108A-OB05-OPT-MIR-04 | Mirror Species | HLATML | HLAT | B108A-OB05-OPT-MIR-04HLAT
W041489 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML6 | B108A-OB05-OPT-MIR-04 | Mirror Species | HPITML | HPIT | B108A-OB05-OPT-MIR-04HPIT
W041491 | MOTION-01 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | ML7 | B108A-FE-DIA-SCRN-01 | Fluorescence Screen Front End 1.5GeV | YML

## MOTION3
Cable ID | Conf. Code | Equip. A | Type A | Chan. A | Equip. B | Type B | Chan. B | Chan. B cleaned | AUX
W041388 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC1 | B108A-FE-OPT-MM-01 | Movable mask unit | XEC | X | B108A-FE-OPT-MM-01X
W041390 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC2 | B108A-FE-OPT-MM-01 | Movable mask unit | YEC | Y | B108A-FE-OPT-MM-01Y
W041392 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC3 | B108A-FE-OPT-MM-02 | Movable mask unit | XEC | X | B108A-FE-OPT-MM-02X
W041394 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC4 | B108A-FE-OPT-MM-02 | Movable mask unit | YEC | Y | B108A-FE-OPT-MM-02Y
W041396 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC5 | B108A-O01-OPT-MIR-01 | Mirror Species | VCEC | VC | B108A-O01-OPT-MIR-01VC
W041398 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC6 | B108A-O01-OPT-MIR-01 | Mirror Species | VREC | VR | B108A-O01-OPT-MIR-01VR
W041400 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC7 | B108A-O01-OPT-MIR-01 | Mirror Species | VLEC | VL | B108A-O01-OPT-MIR-01VL
W041402 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER | EC8 | B108A-O01-OPT-MIR-01 | Mirror Species | HLATEC | HLAT | B108A-O01-OPT-MIR-01HLAT
W041404 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC1 | B108A-O01-OPT-MIR-01 | Mirror Species | HPITEC | HPIT | B108A-O01-OPT-MIR-01HPIT
W041406 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC2 | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PITLEC | M2PIT | B108A-O02-OPT-MONO-01M2PIT
W041408 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC3 | B108A-O02-OPT-MONO-01 | Monochromator Species | GPITLEC | GPIT | B108A-O02-OPT-MONO-01GPIT
W041410 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC4 | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXCEC | GEXC | B108A-O02-OPT-MONO-01GEXC
W041412 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC5 | B108A-O02-OPT-BAFF-01 | Baffle Species | VTEC | VT | B108A-O02-OPT-BAFF-01VT
W041414 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC6 | B108A-O02-OPT-BAFF-01 | Baffle Species | VBEC | VB | B108A-O02-OPT-BAFF-01VB
W041416 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC7 | B108A-O02-OPT-BAFF-01 | Baffle Species | HREC | HR | B108A-O02-OPT-BAFF-01HR
W041418 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE | EC8 | B108A-O02-OPT-BAFF-01 | Baffle Species | HLEC | HL | B108A-O02-OPT-BAFF-01HL
W041420 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC1 | B108A-O03-OPT-BAFF-01 | Baffle Species | VTEC | VT | B108A-O03-OPT-BAFF-01VT
W041422 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC2 | B108A-O03-OPT-BAFF-01 | Baffle Species | VBEC | VB | B108A-O03-OPT-BAFF-01VB
W041424 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC3 | B108A-O03-OPT-BAFF-01 | Baffle Species | HLEC | HL | B108A-O03-OPT-BAFF-01HL
W041426 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC4 | B108A-O03-OPT-BAFF-01 | Baffle Species | HREC | HR | B108A-O03-OPT-BAFF-01HR
W041428 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC5 | B108A-O03-OPT-MIR-03 | Mirror Species | VCEC | VC | B108A-O03-OPT-MIR-03VC
W041430 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC6 | B108A-O03-OPT-MIR-03 | Mirror Species | VREC | VR | B108A-O03-OPT-MIR-03VR
W041432 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC7 | B108A-O03-OPT-MIR-03 | Mirror Species | VLEC | VL | B108A-O03-OPT-MIR-03VL
W041434 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE | EC8 | B108A-O03-OPT-MIR-03 | Mirror Species | HLATEC | HLAT | B108A-O03-OPT-MIR-03HLAT
W041436 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | EC1 | B108A-O03-OPT-MIR-03 | Mirror Species | HPITEC | HPIT | B108A-O03-OPT-MIR-03HPIT
W041438 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | EC2 | B108A-O03-OPT-MIR-03 | Mirror Species | MCHGEC | MCHG | B108A-O03-OPT-MIR-03MCHG
W041440 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | EC3 | B108A-OA04-DIA-BPM-01 | BPM Species | VEC | V | B108A-OA04-DIA-BPM-01V
W041442 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | EC4 | B108A-OB04-DIA-BPM-01 | BPM Species | VEC | V | B108A-OB04-DIA-BPM-01V
W041444 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | EC5 | B108A-OA04-OPT-SLIT-01 | Slits Species | VEC | V | B108A-OA04-OPT-SLIT-01V
W041448 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE | EC8 | B108A-OA04-OPT-SLIT-01 | Slits Species | HZEC | HZ | B108A-OA04-OPT-SLIT-01HZ
W041450 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | EC1 | B108A-OB04-OPT-SLIT-01 | Slits Species | VEC | V | B108A-OB04-OPT-SLIT-01V
W041454 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | EC4 | B108A-OB04-OPT-SLIT-01 | Slits Species | HZEC | HZ | B108A-OB04-OPT-SLIT-01HZ
W041456 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | EC5 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VTEC | VT | B108A-OA05-OPT-BAFF-01VT
W041458 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | EC6 | B108A-OA05-OPT-BAFF-01 | Baffle Species | VBEC | VB | B108A-OA05-OPT-BAFF-01VB
W041460 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | EC7 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HLEC | HL | B108A-OA05-OPT-BAFF-01HL
W041462 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE | EC8 | B108A-OA05-OPT-BAFF-01 | Baffle Species | HREC | HR | B108A-OA05-OPT-BAFF-01HR
W041464 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC1 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VTEC | VT | B108A-OB05-OPT-BAFF-01VT
W041466 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC2 | B108A-OB05-OPT-BAFF-01 | Baffle Species | VBEC | VB | B108A-OB05-OPT-BAFF-01VB
W041468 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC3 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HLEC | HL | B108A-OB05-OPT-BAFF-01HL
W041470 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC4 | B108A-OB05-OPT-BAFF-01 | Baffle Species | HREC | HR | B108A-OB05-OPT-BAFF-01HR
W041472 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC5 | B108A-OA05-OPT-MIR-04 | Mirror Species | VCEC | VC | B108A-OA05-OPT-MIR-04VC
W041474 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC6 | B108A-OA05-OPT-MIR-04 | Mirror Species | VREC | VR | B108A-OA05-OPT-MIR-04VR
W041476 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC7 | B108A-OA05-OPT-MIR-04 | Mirror Species | VLEC | VL | B108A-OA05-OPT-MIR-04VL
W041478 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE | EC8 | B108A-OA05-OPT-MIR-04 | Mirror Species | HLATEC | HLAT | B108A-OA05-OPT-MIR-04HLAT
W041480 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC1 | B108A-OA05-OPT-MIR-04 | Mirror Species | HPITEC | HPIT | B108A-OA05-OPT-MIR-04HPIT
W041482 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC2 | B108A-OB05-OPT-MIR-04 | Mirror Species | VCEC | VC | B108A-OB05-OPT-MIR-04VC
W041484 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC3 | B108A-OB05-OPT-MIR-04 | Mirror Species | VREC | VR | B108A-OB05-OPT-MIR-04VR
W041486 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC4 | B108A-OB05-OPT-MIR-04 | Mirror Species | VLEC | VL | B108A-OB05-OPT-MIR-04VL
W041488 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC5 | B108A-OB05-OPT-MIR-04 | Mirror Species | HLATEC | HLAT | B108A-OB05-OPT-MIR-04HLAT
W041490 | MOTION-03 | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE | EC6 | B108A-OB05-OPT-MIR-04 | Mirror Species | HPITEC | HPIT | B108A-OB05-OPT-MIR-04HPIT

## DISABLE
Cable ID | Conf. Code | Equip. A | Type A | Chan. A | Equip. B | Type B | Chan. B | Chan. B cleaned | AUX
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-FE-OPT-MM-01 | Movable mask unit | XDIS | X | B108A-FE-OPT-MM-01X
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-FE-OPT-MM-01 | Movable mask unit | YDIS | Y | B108A-FE-OPT-MM-01Y
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-FE-OPT-MM-02 | Movable mask unit | XDIS | X | B108A-FE-OPT-MM-02X
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-FE-OPT-MM-02 | Movable mask unit | YDIS | Y | B108A-FE-OPT-MM-02Y
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-O01-OPT-MIR-01 | Mirror Species | VCDIS | VC | B108A-O01-OPT-MIR-01VC
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-O01-OPT-MIR-01 | Mirror Species | VRDIS | VR | B108A-O01-OPT-MIR-01VR
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-O01-OPT-MIR-01 | Mirror Species | VLDIS | VL | B108A-O01-OPT-MIR-01VL
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-01 | ESRF IcePAP MASTER |  | B108A-O01-OPT-MIR-01 | Mirror Species | HLATDIS | HLAT | B108A-O01-OPT-MIR-01HLAT
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O01-OPT-MIR-01 | Mirror Species | HPITDIS | HPIT | B108A-O01-OPT-MIR-01HPIT
- | FLAT09R-01 | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O02-OPT-MONO-01 | Monochromator Species | M2PITDIS | M2PIT | B108A-O02-OPT-MONO-01M2PIT
- |  | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O02-OPT-MONO-01 | Monochromator Species | GPITDIS | GPIT | B108A-O02-OPT-MONO-01GPIT
- |  | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O02-OPT-MONO-01 | Monochromator Species | GEXCDIS | GEXC | B108A-O02-OPT-MONO-01GEXC
- |  | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O02-OPT-BAFF-01 | Baffle Species | VTDIS | VT | B108A-O02-OPT-BAFF-01VT
- |  | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O02-OPT-BAFF-01 | Baffle Species | VBDIS | VB | B108A-O02-OPT-BAFF-01VB
- |  | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O02-OPT-BAFF-01 | Baffle Species | HRDIS | HR | B108A-O02-OPT-BAFF-01HR
- |  | B108A-D100830-CAB04-CTL-IPAP-02 | ESRF IcePAP SLAVE |  | B108A-O02-OPT-BAFF-01 | Baffle Species | HLDIS | HL | B108A-O02-OPT-BAFF-01HL
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-BAFF-01 | Baffle Species | VTDIS | VT | B108A-O03-OPT-BAFF-01VT
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-BAFF-01 | Baffle Species | VBDIS | VB | B108A-O03-OPT-BAFF-01VB
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-BAFF-01 | Baffle Species | HLDIS | HL | B108A-O03-OPT-BAFF-01HL
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-BAFF-01 | Baffle Species | HRDIS | HR | B108A-O03-OPT-BAFF-01HR
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-MIR-03 | Mirror Species | VCDIS | VC | B108A-O03-OPT-MIR-03VC
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-MIR-03 | Mirror Species | VRDIS | VR | B108A-O03-OPT-MIR-03VR
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-MIR-03 | Mirror Species | VLDIS | VL | B108A-O03-OPT-MIR-03VL
- |  | B108A-D100830-CAB04-CTL-IPAP-03 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-MIR-03 | Mirror Species | HLATDIS | HLAT | B108A-O03-OPT-MIR-03HLAT
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-MIR-03 | Mirror Species | HPITDIS | HPIT | B108A-O03-OPT-MIR-03HPIT
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-O03-OPT-MIR-03 | Mirror Species | MCHGDIS | MCHG | B108A-O03-OPT-MIR-03MCHG
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-OA04-DIA-BPM-01 | BPM Species | VDIS | V | B108A-OA04-DIA-BPM-01V
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-OB04-DIA-BPM-01 | BPM Species | VDIS | V | B108A-OB04-DIA-BPM-01V
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-OA04-OPT-SLIT-01 | Slits Species | VDIS | V | B108A-OA04-OPT-SLIT-01V
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-OA04-OPT-SLIT-01 | Slits Species | HDARDIS | HDAR | B108A-OA04-OPT-SLIT-01HDAR
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-OA04-OPT-SLIT-01 | Slits Species | HDALDIS | HDAL | B108A-OA04-OPT-SLIT-01HDAL
- |  | B108A-D100830-CAB04-CTL-IPAP-04 | ESRF IcePAP SLAVE |  | B108A-OA04-OPT-SLIT-01 | Slits Species | HZDIS | HZ | B108A-OA04-OPT-SLIT-01HZ
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OB04-OPT-SLIT-01 | Slits Species | VDIS | V | B108A-OB04-OPT-SLIT-01V
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OB04-OPT-SLIT-01 | Slits Species | HDARDIS | HDAR | B108A-OB04-OPT-SLIT-01HDAR
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OB04-OPT-SLIT-01 | Slits Species | HDALDIS | HDAL | B108A-OB04-OPT-SLIT-01HDAL
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OB04-OPT-SLIT-01 | Slits Species | HZDIS | HZ | B108A-OB04-OPT-SLIT-01HZ
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-BAFF-01 | Baffle Species | VTDIS | VT | B108A-OA05-OPT-BAFF-01VT
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-BAFF-01 | Baffle Species | VBDIS | VB | B108A-OA05-OPT-BAFF-01VB
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-BAFF-01 | Baffle Species | HLDIS | HL | B108A-OA05-OPT-BAFF-01HL
- |  | B108A-D100830-CAB04-CTL-IPAP-05 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-BAFF-01 | Baffle Species | HRDIS | HR | B108A-OA05-OPT-BAFF-01HR
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-BAFF-01 | Baffle Species | VTDIS | VT | B108A-OB05-OPT-BAFF-01VT
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-BAFF-01 | Baffle Species | VBDIS | VB | B108A-OB05-OPT-BAFF-01VB
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-BAFF-01 | Baffle Species | HLDIS | HL | B108A-OB05-OPT-BAFF-01HL
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-BAFF-01 | Baffle Species | HRDIS | HR | B108A-OB05-OPT-BAFF-01HR
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-MIR-04 | Mirror Species | VCDIS | VC | B108A-OA05-OPT-MIR-04VC
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-MIR-04 | Mirror Species | VRDIS | VR | B108A-OA05-OPT-MIR-04VR
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-MIR-04 | Mirror Species | VLDIS | VL | B108A-OA05-OPT-MIR-04VL
- |  | B108A-D100830-CAB04-CTL-IPAP-06 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-MIR-04 | Mirror Species | HLATDIS | HLAT | B108A-OA05-OPT-MIR-04HLAT
- |  | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE |  | B108A-OA05-OPT-MIR-04 | Mirror Species | HPITDIS | HPIT | B108A-OA05-OPT-MIR-04HPIT
- |  | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-MIR-04 | Mirror Species | VCDIS | VC | B108A-OB05-OPT-MIR-04VC
|  | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-MIR-04 | Mirror Species | VRDIS | VR | B108A-OB05-OPT-MIR-04VR
|  | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-MIR-04 | Mirror Species | VLDIS | VL | B108A-OB05-OPT-MIR-04VL
|  | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-MIR-04 | Mirror Species | HLATDIS | HLAT | B108A-OB05-OPT-MIR-04HLAT
|  | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE |  | B108A-OB05-OPT-MIR-04 | Mirror Species | HPITDIS | HPIT | B108A-OB05-OPT-MIR-04HPIT
|  | B108A-D100830-CAB04-CTL-IPAP-07 | ESRF IcePAP SLAVE |  | B108A-FE-DIA-SCRN-01 | Fluorescence Screen Front End 1.5GeV | YDIS | Y | B108A-FE-DIA-SCRN-01Y

## MOTORS
Number of motors: | 100
Motor model and wiring (if many) | Vnominal | Iphase(A) | Rphase(Ohm) | Lphase(mH) | full steps | NVOLT | IVOLT | NCURR | ICURR | MOTPOLES | ANTURN | ANSTEP | DEFVEL | DEFACCT | MREGMODE | CURRGAIN | HOMEVEL | ABSMODE | SSIDBITS | SSICODE | SSICLOCK | SSIDELAY | SSISTATUS | SSISTRTB | SSIMSKNB | SSIMSKMSB | SSILDC | SSIALDC | SSIEWCHK | SSIEWPOL | SSIEWPOS | SSIEECHK | SSIEEPOL | SSIEEPOS
Stoegra SM 56.1.18.J1 | 60 | 2 | 4.75 | 9 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Stoegra SM 56.2.18.J3 | 60 | 5.9 | 1.2 | 2.6 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Stoegra SM 56.3.18.J4.6 | 60 | 7.5 | 0.72 | 1.2 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 19.200.0.6 parallel | 25 | 0.8 | 1.9 | 0.5 | 200 | 25 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 25.200.1.2 parallel | 25 | 1.7 | 0.95 | 0.4 | 200 | 25 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 32.200.1.2 parallel | 42 | 1.7 | 1.25 | 1.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 32.200.2.5 parallel | 25 | 3.5 | 0.3 | 0.3 | 200 | 25 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 41.500.0.6 parallel | 42 | 0.8 | 5.1 | 7.6 | 500 | 42 | 12 | 0.8 | 50 | 125 | 1 | 1000 | 2500 | 0.3 | CURR | LOW | 1250
Phytron ZSS 42.200.1.2 parallel | 42 | 1.7 | 1.7 | 3 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 43.500.2.5 parallel | 42 | 3.5 | 0.5 | 1.2 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS.52.200.2.5 parallel | 48 | 3.5 | 0.6 | 1.6 | 200 | 48 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS.52.200.2.5 series | 48 | 1.8 | 2.4 | 6.4 | 200 | 48 | 12 | 1.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 57.200.1.2 parallel | 48 | 1.7 | 3.9 | 7.8 | 200 | 48 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSS 57.200.2.5 parallel | 48 | 3.5 | 0.8 | 2.4 | 200 | 48 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 57/3.200.6.5 parallel | 48 | 7.5 | 0.4 | 1.2 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 88/3.200.8 parallel | 60 | 7.5 | 0.14 | 2.6 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 87/2.200.6.5 parallel | 60 | 7.5 | 0.25 | 1.5 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 87/2.200.8.4 parallel | 60 | 7.5 | 0.15 | 1 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron ZSH 88/2.200.8 parallel | 60 | 7.5 | 0.1 | 1.5 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 19.200.0.6 parallel | 25 | 0.8 | 2.1 | 0.55 | 200 | 25 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 25.200.1.2 parallel | 25 | 1.7 | 0.95 | 0.4 | 200 | 25 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 32.200.1.2 parallel | 42 | 1.7 | 1.25 | 1.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 33.200.1.2 parallel | 42 | 1.7 | 1.9 | 2.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 42.200.1.2 parallel | 42 | 1.7 | 1.7 | 3 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 42.500.2.5 parallel | 25 | 3.5 | 0.34 | 0.7 | 500 | 25 | 12 | 3.5 | 50 | 125 | 1 | 1000 | 2500 | 0.3 | CURR | LOW | 1250
Phytron VSS 43.200.1.2 parallel | 42 | 1.7 | 1.3 | 5.2 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 43.200.2.5 parallel | 42 | 3.5 | 0.5 | 1.2 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 52.200.1.2 parallel | 42 | 1.7 | 2.65 | 7 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 52.200.2.5 parallel | 42 | 3.5 | 0.6 | 1.6 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 52.200.5 parallel | 25 | 7 | 0.17 | 0.4 | 200 | 25 | 20 | 7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 57.200.1.2 parallel | 42 | 1.7 | 3.9 | 9.5 | 200 | 42 | 12 | 1.7 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Phytron VSS 57.200.2.5 parallel | 42 | 3.5 | 0.8 | 2.4 | 200 | 42 | 12 | 3.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Superior Slo-syn KML062F13 | 48 | 7.5 | 0.2 | 0.85 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Superior Slo-syn KML061F05 | 48 | 3.8 | 0.85 | 4.56 | 200 | 48 | 12 | 3.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Superior Slo-syn KML061F02 | 48 | 1.5 | 4.94 | 30.1 | 200 | 48 | 12 | 1.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK223PB series | 48 | 0.9 | 5.6 | 4 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK243-01 series | 48 | 0.9 | 8.4 | 10 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK243-02 series | 48 | 0.4 | 48 | 60 | 200 | 48 | 12 | 0.4 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK244-01 series | 48 | 1.2 | 6.6 | 12.8 | 200 | 48 | 12 | 1.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK244-01 parallel | 48 | 2.4 | 1.8 | 3.2 | 200 | 48 | 12 | 2.4 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK245-01 series | 48 | 1.2 | 6.6 | 11.2 | 200 | 48 | 12 | 1.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK245D | 48 | 2.1 | 2.1 | 3.1 | 200 | 48 | 12 | 2.1 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK245MD15 | 48 | 2.1 | 2.12 | 5.1 | 400 | 48 | 12 | 2.1 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK245M-02 series | 48 | 0.8 | 15 | 39.6 | 400 | 48 | 12 | 0.8 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK245M-02 half coil | 48 | 1.1 | 7.5 | 9.9 | 400 | 48 | 12 | 1.1 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK264-02 series | 48 | 2 | 2.8 | 5.6 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264-E2.0 series | 48 | 2 | 2.8 | 5.6 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264M-E2.0 parallel | 48 | 3.9 | 0.7 | 1.7 | 400 | 48 | 12 | 3.9 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK264BE-SGxx series | 48 | 2 | 2.8 | 5.6 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264BE-SGxx parallel | 48 | 3.9 | 0.7 | 1.4 | 200 | 48 | 12 | 3.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK264JD | 48 | 3.9 | 0.73 | 1.8 | 200 | 48 | 12 | 3.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK266-02 series | 48 | 2 | 3.6 | 10 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK266D14 | 48 | 2 | 3.6 | 10 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK267J series | 48 | 2 | 4.8 | 14.2 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK267JD series | 48 | 3.9 | 1.2 | 3.54 | 200 | 48 | 12 | 3.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268-E2.0 series | 48 | 2 | 4.5 | 14.4 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268M-E2.0 parallel | 48 | 3.9 | 1.13 | 4.8 | 400 | 48 | 12 | 3.9 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK268-01 series | 48 | 1 | 17.2 | 56 | 200 | 48 | 12 | 1 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268-02 series | 48 | 2 | 4.5 | 14.4 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268D14 | 48 | 2 | 4.5 | 14.4 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK268M-E2.0 series | 48 | 2 | 4.5 | 19.2 | 400 | 48 | 12 | 2 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK268M-01 half coil | 48 | 1.4 | 8.6 | 19.4 | 400 | 48 | 12 | 1.4 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK268M-02 series | 48 | 2 | 4.5 | 19.2 | 400 | 48 | 12 | 2 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PK268-03 series | 48 | 2.9 | 2 | 6.4 | 200 | 48 | 12 | 2.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK269JD | 48 | 3.9 | 1.49 | 5.7 | 200 | 48 | 12 | 3.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296-E4.5 series | 48 | 4.5 | 0.96 | 6 | 200 | 48 | 12 | 4.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296-E4.5 parallel | 48 | 7.5 | 0.24 | 1.5 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296AE-SGxx series | 48 | 2.9 | 0.96 | 6 | 200 | 48 | 12 | 2.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK296AE-SGxx parallel | 48 | 5.9 | 0.24 | 1.5 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK299-E4.5 series | 48 | 4.5 | 0.96 | 6 | 200 | 48 | 12 | 4.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PK299-E4.5 parallel | 48 | 7.5 | 0.24 | 1.5 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Oriental PKP264MU20B-L half coil | 48 | 2.8 | 1.45 |  | 400 | 48 | 12 | 2.8 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Oriental PKP268MD28B-L | 48 | 3.9 | 1.23 | 6 | 400 | 48 | 12 | 3.9 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 500
Faulhaber Precistep AM1524-2R-A-0,25-12,5 | 20 | 0.4 | 12.5 | 6.3 | 24 | 20 | 12 | 0.4 | 50 | 6 | 1 | 48 | 120 | 0.3 | CURR | LOW | 60
Faulhaber Precistep AM1524-2R-A-0,45-3,6 | 20 | 0.6 | 3.6 | 1.9 | 24 | 20 | 12 | 0.6 | 50 | 6 | 1 | 48 | 120 | 0.3 | CURR | LOW | 60
McLennan 23HT18C230 parallel | 48 | 5.9 | 0.38 | 1.2 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HS-309 series | 48 | 3.4 | 1.48 | 2.92 | 200 | 48 | 12 | 3.4 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HS-309 parallel | 48 | 6.6 | 0.37 | 0.73 | 200 | 48 | 20 | 6.6 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 17HS-240E | 48 | 3.2 | 0.72 | 0.83 | 200 | 48 | 12 | 3.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HS-108 Mk2 | 48 | 5.5 | 0.37 | 0.59 | 200 | 48 | 20 | 5.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HSX-206 parallel | 48 | 5.9 | 0.35 | 0.9 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
McLennan 23HSX-306E parallel | 48 | 5.9 | 0.55 | 1.7 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Berger Lahr VRDM 268/50L3G8A | 48 | 4.3 | 2.2 | 20 | 200 | 48 | 12 | 4.3 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
PowerMaxII P21NRXD-LNF-NS-00 | 48 | 1.1 | 5 | 10 | 200 | 48 | 12 | 1.1 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear Engineering 3518X-04D-0 | 48 | 0.6 | 3.8 | 2.7 | 200 | 48 | 12 | 0.6 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear Engineering 5718X-18DE-01 series | 48 | 4.2 | 0.8 | 1.4 | 200 | 48 | 12 | 4.2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear Engineering 5718X-18DE-01 parallel | 48 | 7.5 | 0.2 | 0.35 | 200 | 48 | 20 | 7.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear engineering 41185-04PD-23RO parallel | 48 | 1.9 | 2.5 | 3.1 | 200 | 48 | 12 | 1.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear engineering 4118S-04s series | 48 | 0.9 | 9.9 | 12.5 | 200 | 48 | 12 | 0.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Linear engineering 4118S-04s parallel | 48 | 1.9 | 2.5 | 3.1 | 200 | 48 | 12 | 1.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotec ST5918L3008 series | 48 | 2.9 | 2 | 8.8 | 200 | 48 | 12 | 2.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotec ST5918L3008 parallel | 48 | 5.9 | 0.5 | 2.2 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotec L4118S1404-M6X1 | 48 | 2 | 2 | 2.8 | 200 | 48 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotech AS4118L1804-E | 48 | 2.5 | 1.75 | 3.3 | 200 | 48 | 12 | 2.5 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Nanotech AS5918L4204-E | 48 | 5.9 | 0.5 | 1.9 | 200 | 48 | 20 | 5.9 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Standa 8cma20-8/15 | 24 | 0.8 | 5.4 | 1.4 | 200 | 24 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Standa PG15/50 | 24 | 0.7 | 10 | 3 | 20 | 24 | 12 | 0.7 | 50 | 5 | 1 | 40 | 100 | 0.3 | CURR | LOW | 50
Wantai 42BYGHM809 | 48 | 2.4 | 1.8 | 2.8 | 400 | 48 | 12 | 2.4 | 50 | 100 | 1 | 800 | 2000 | 0.3 | CURR | LOW | 1000
Fulling 28STH45B2-096 series | 24 | 0.8 | 6.8 | 4.9 | 200 | 24 | 12 | 0.8 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Trinamic Qmot QSH-4128-35-10-27 | 24 | 2 | 5.3 | 6.6 | 200 | 24 | 12 | 2 | 50 | 50 | 1 | 400 | 1000 | 0.3 | CURR | LOW | 500
Newport TraXXppd | 24 | 0.4 | 12.5 | 5.5 | 20 | 24 | 12 | 0.4 | 50 | 5 | 1 | 40 | 100 | 0.3 | CURR | LOW | 50
Renishaw Resolute 32 bits | [17 empty] | BISSC | 0 | BINARY | 1.25MHz | 0 | S | 0 | 0 | LSB | 0 | AUTO | YES | LOW | 6 | YES | LOW | 7
Renishaw Resolute 26 bits | [17 empty] | BISSC | 26 | BINARY | 1.25MHz | 0 | S | 0 | 0 | LSB | 0 | AUTO | YES | LOW | 6 | YES | LOW | 7
Renishaw Resolute 36 bits | [17 empty] | BISSC | 1 | BINARY | 1.25MHz | 0 | S | 0 | 5 | LSB | 0 | AUTO | YES | LOW | 6 | YES | LOW | 7
RLS LMA10DCxxxxxxxx | [17 empty] | BISSC | 26 | BINARY | 1.25MHz | 0 | S | 1 | 0 | LSB | 0 | AUTO | NO | LOW | 6 | YES | HIGH | 7
Renishaw Resolute 32 bits MSBMSK 1 (31b) | [17 empty] | BISSC | 0 | BINARY | 1.25MHz | 0 | S | 0 | 1 | LSB | 0 | AUTO | YES | LOW | 6 | YES | LOW | 7